## CF&PCF-190,1200,1500
BÁO CÁO GIÁM ĐỊNH KỸ THUẬT
TECHNICAL REPORT
SỐ/ No.: |  |  |  |  | NGÀY RA BÁO CÁO/ Issued Date:
KÍNH GỞI/ Report to: |  |  | Quý Khách Hàng/ Valued Customer
BỆNH VIỆN/ Name Of User: | [8 empty] | CV-170
KIỂU MÁY/ Name Of Model: |  |  | CV-170 |  | SỐ MÁY/ Serial No: |  |  | 7544794 | CHO RAY HOSPITAL
HƯ HỎNG TỪ KHÁCH HÀNG/ Customer Complaints:
1. Rò rỉ nút số 1/ Leaking button number 1 2.
ĐÁNH GIÁ TỪ OLYMPUS/ Olympus findings
Bộ Phận/ Section |  | Đánh Giá Hư Hỏng/ Confirmed Defects |  |  |  |  | Hình Ảnh/ Photos
Bộ phận uốn cong Bending Section |  | Hiện Trạng/ Phenomenon
1. Keo dán cao su di động bị bong tróc The cement of the A-rubber is detached
2. Bộ phận uốn cong hướng xuống không đủ (160°)/ The Down direction of Bending section is not enough (160°) Tiêu chuẩn là 180° +10/-10/ Standard angle is 180° +10/-10
Đánh giá/ Judgement
Yêu cầu sửa chữa/ Repair required
Bộ phận điều khiển Control Unit |  | Hiện Trạng/ Phenomenon
Nút điều khiển số 1 bị rách/ The Switch No.1 is torn
Đánh giá/ Judgement
Yêu cầu sửa chữa/ Repair required
Dây dẫn chung Universal Cord |  | Hiện Trạng/ Phenomenon
Dây dẫn chung bị nhăn/ The Universal cord is wrinkled
Đánh giá/ Judgement
Khuyến nghị sửa chữa/ Repair recommended
GHI CHÚ/ Remark
NGUYÊN NHÂN HƯ HỎNG/ Reason Of Damages
- Các hư hỏng do hao mòn trong quá trình sử dụng/ The damages due to worn and tear during use.
GIẢI PHÁP/ Solutions
- Thay thế linh kiện như danh mục chuẩn đính kèm. Tuy nhiên nếu khách hàng không đồng ý sửa chữa lớn, chúng tôi chỉ thay thế linh kiện như danh mục khuyến nghị kèm theo. Chúng tôi chỉ bảo hành cho các linh kiện thay thế này.
Replace new spare parts as standard attachment list. However, if the customer does not agree to Major repair, we will replace the new spare parts as recommended attachment list. We only warrant replaced parts.
HÀNH ĐỘNG NGĂN NGỪA/ Recommended Actions To Prevent Sử dụng, bảo quản và vệ sinh tiệt trùng thiết bị theo đúng hướng dẫn của Olympus. Use, maintain and reprocess the equipment following Olympus’s instructions.
- Sử dụng, bảo quản và vệ sinh tiệt trùng thiết bị theo đúng hướng dẫn của Olympus. - Use, maintain and reprocess the equipment following Olympus’s instructions.
THÔNG TIN SỬA CHỮA/ Repair Information
| SỬA CHỮA THEO NHU CẦU On-demand Repair |  |  | TRONG BẢO HÀNH Under Warranty |  |  |  | BẢO HÀNH MUA MỚI Product Warranty
BẢO HÀNH SỬA CHỮA Repair Warranty
KHÔNG ĐƯỢC BẢO HÀNH No Warranty
HẾT BẢO HÀNH Out Of Warranty
| HỢP ĐỒNG BẢO TRÌ Service Contract |  |  | THUỘC ĐIỀU KHOẢN HỢP ĐỒNG No-charge
NGOÀI ĐIỀU KHOẢN HỢP ĐỒNG Charge
THUỘC MỘT PHẦN ĐIỀU KHOẢN HỢP ĐỒNG Particular Charge
ĐẠI DIỆN CÔNG TY OLYMPUS | [5 empty] | ĐẠI DIỆN KHÁCH HÀNG
Representative of Olympus Company | [5 empty] | Representative of Customer
NGƯỜI KIỂM TRA/ Inspected By
Nguyễn Hoàng Vũ
South Region Technical Service Engineer
NGƯỜI DUYỆT/ Approved By
Phạm Trung Vũ
South Region Technical Service Assistant Manager
OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD. (OVN)
Office and Service Center in Hanoi: Unit 305- 307, 3rd Floor, V Tower, 649 Kim Ma Street, Giang Vo Ward, HaNoi City, Vietnam - Tel: (84-24) 3766 5831/32/34
Office In HoChiMinh City: 7th Floor, Viettel Building, 285 Cach Mang Thang Tam, Hoa Hung Ward, HoChiMinh City, Vietnam - Tel: (84-28) 7301 9919
Service Center In HoChiMinh City: 2nd Floor, 86 Tan Da, Cho Lon Ward, HoChiMinh City, Vietnam - Tel: (84-28) 7301 9919
Endoscope Connector
Universal Cord
Control Section
Insertion Tube
Bending Section
Distal End
Scope Structure

## DATA
Model | Serial No. | Enduser
CH-S400-XZ-EB | 7823404 | PEOPLE'S HOSPITAL 115
CF-H170I | 2610743 | LUONG TAI DISTRICT GENERAL HOSPITAL
GIF-H170 | 2828926 | CHO RAY HOSPITAL
CF-H170I | 2711274 | GIA DINH PEOPLE HOSPITAL
GIF-V70-PAL | 2411510 | HOC MON REGION GENERAL HOSPITAL
BF-1T150 | 2140511 | CHO RAY HOSPITAL
GIF-HQ190 | 2746560 | TRUNG VUONG HOSPITAL
GIF-Q150 | 2519789 | TRUNG VUONG HOSPITAL
LF-TP(STERRAD) | 2739052 | CHO RAY HOSPITAL
CV-170 | 7544794 | CHO RAY HOSPITAL
CV-170 | 7883541 | CHO RAY HOSPITAL
GIF-H170 | 2512434 | CHO RAY HOSPITAL
GIF-H170 | 2512427 | CHO RAY HOSPITAL
GIF-H170 | 2827849 | CHO RAY HOSPITAL
CF-H170I | 2822235 | CHO RAY HOSPITAL
GIF-H170 | 2512426 | CHO RAY HOSPITAL
ESG-400 | B002949 | PEOPLE'S HOSPITAL 115
CF-HQ190I | 2510963 | CHO RAY HOSPITAL
GIF-HQ190 | 2519678 | CHO RAY HOSPITAL
TJF-150 | 2000470 | FRIENDSHIP NGHE AN GENERAL HOSPITAL
GIF-H170 | 2827853 | CHO RAY HOSPITAL
LF-GP | 2118931 | GIA DINH PEOPLE HOSPITAL
GIF-Q150 | 2901689 | LANG SON PROVINCIAL GENERAL HOSPITAL
BF-1TQ170 | 2700370 | PHAM NGOC THACH HOSPITAL
GIF-Q150 | 2620891 | SAINT PAUL GENERAL HOSPITAL
GIF-Q180 | 2102809 | SAINT PAUL GENERAL HOSPITAL
CF-Q180AL | 2208765 | QUANG NAM CENTRAL GENERAL HOSPITAL
CV-170 | 7544789 | CHO RAY HOSPITAL
CV-170 | 7544788 | CHO RAY HOSPITAL
CHF-V(STERRAD NX) | 2651152 | TRUNG VUONG HOSPITAL
CF-Q180AL | 2008485 | SAINT PAUL GENERAL HOSPITAL
CF-H180AI | 2202768 | THONG NHAT HOSPITAL
OTV-S190 | 7703923 | BINH DAN HOSPITAL
SPL-G | CG0677 | BINH DAN HOSPITAL
CF-H190I | 2621702 | SAINT PAUL GENERAL HOSPITAL
GIF-H170 | 2512432 | CHO RAY HOSPITAL
GIF-H170 | 2929881 | THU DUC CITY HOSPITAL
TJF-Q180V | 2507310 | CHO RAY HOSPITAL
OTV-S190 | 7703922 | BINH DAN HOSPITAL
PCF-PH190I | 2840236 | PEDIATRIC HOSPITAL NO. 1
GIF-Q180 | 2809016 | BACH MAI HOSPITAL
GIF-Q150 | 2104234 | UNIVERSITY MEDICAL CENTER HCMC 2
GIF-Q150 | 2821675 | UNIVERSITY MEDICAL CENTER HCMC 2
GIF-HQ190 | 2746172 | HUNG HA GENERAL HOSPITAL
GIF-Q150 | 2306558 | UNIVERSITY MEDICAL CENTER HCMC
CF-H170I | 2610931 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
CF-H170I | 2711252 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
URF-V | 2350910 | DA NANG C HOSPITAL
GIF-Q150 | 2307385 | DONG NAI GENERAL HOSPITAL
GIF-Q180 | 2304322 | DONG NAI GENERAL HOSPITAL
GIF-H170 | 2512695 | E HOSPITAL
UHI-4 | 7740062 | QUANG NGAI OBSTETRICS AND PEDIATRICS HOSPITAL
TJF-150 | 2611869 | HAI PHONG INTERNATIONAL HOSPITAL
GIF-Q150 | 2922252 | GIA DINH PEOPLE HOSPITAL
GIF-H170 | 2930352 | DA NANG HOSPITAL
CF-Q150I | 2100743 | NATIONAL TRANSPORT HOSPITAL
BF-1T150 | 2800112 | CHO RAY HOSPITAL
CV-170 | 7986655 | DA NANG HOSPITAL
GIF-Q150 | 2519528 | E HOSPITAL
CF-H170I | 2500348 | HONG NGOC GENERAL HOSPITAL
WA2T412A | 808713 | UNIVERSITY MEDICAL CENTER HCMC
GIF-H170 | 2613244 | BINH DINH PROVINCE GENERAL HOSPITAL
CF-H170I | 2821840 | UNIVERSITY MEDICAL CENTER HCMC 2
ENF-V4 | 2900371 | 108 MILITARY CENTRAL HOSPITAL
CHF-V(STERRAD NX) | 2062477 | TRUNG VUONG HOSPITAL
GIF-H170 | 2614316 | SHINGMARK MEDICAL UNIVERSITY HOSPITAL
GIF-H170 | 2614126 | THU DUC CITY HOSPITAL
GIF-H170 | 2827917 | CU CHI REGIONAL GENERAL HOSPITAL
GIF-H170 | 2512445 | GOLD STAR GENERAL HOSPITAL
GIF-H170 | 2033241 | 110 MILITARY HOSPITAL
GIF-Q150 | 2620629 | HAI DUONG GENERAL HOSPITAL
GIF-H170 | 2614074 | YEN LAC DISTRICT HOSPITAL
GIF-H170 | 2715497 | ANH KHOA MEDICAL EQUIPMENT Co., LTD
GIF-H170 | 2501737 | THANH NHAN HOSPITAL
CF-Q150I | 2100607 | THANH HA DISTRICT HOSPITAL
BF-P150(STERRAD NX) | 2651116 | NATIONAL LUNG HOSPITAL
BF-XP190 | 2811286 | PEDIATRIC HOSPITAL NO. 1
OFP-2 | 220101873 | HOAN MY SAIGON HOSPITAL
GIF-H170 | 2614347 | VICTORIA CLINIC
GIF-H170 | 2501639 | TIEN GIANG CENTRAL GENERAL HOSPITAL
GIF-Q180 | 2707629 | BINH PHUOC PROVINCE GENERAL HOSPITAL
GIF-H170 | 2931166 | HA TINH GENERAL HOSPITAL
GIF-Q180 | 2900838 | DONG NAI GENERAL HOSPITAL
GIF-H170 | 2031490 | HOAN MY THU DUC INTERNATIONAL GENERAL HOSPITAL
ESG-400 | B002051 | BINH DAN HOSPITAL
ESG-400 | B007018 | BINH DAN HOSPITAL
BF-1T150 | 2140444 | PHAM NGOC THACH HOSPITAL
OTV-S7PROH-HD-12E | 7029243 | AN BINH HOSPITAL
CYF-5A(STERRAD NX) | 2963161 | BINH DAN HOSPITAL
URF-V(STERRAD NX) | 2973527 | BINH DAN HOSPITAL
TJF-Q180V | 2404798 | BECAMEX INTERNATIONAL HOSPITAL
BF-P180(STERRAD NX) | 2242963 | K74 HOSPITAL
CF-Q150I | 2601854 | HAI DUONG GENERAL HOSPITAL
BF-1T150 | 2140610 | NATIONAL LUNG HOSPITAL
ESG-400 | 15268W140008 | BINH DINH PROVINCE GENERAL HOSPITAL
CV-180 | 7103257 | PHAM NGOC THACH HOSPITAL OF QUANG NAM
GIF-XP170N | 2931834 | AN SINH HOSPITAL
WA4KL130 | 804135 | THONG NHAT HOSPITAL
URF-V | 2140613 | UNIVERSITY MEDICAL CENTER HCMC
GIF-H190 | 2959906 | BAO SON PRIVATE HOSPITAL
GIF-Q150 | 2922250 | GIA DINH PEOPLE HOSPITAL
GIF-Q150 | 2922249 | GIA DINH PEOPLE HOSPITAL
GIF-Q180 | 2304310 | DONG NAI GENERAL HOSPITAL
CF-H180AI | 2901377 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
OTV-S190 | 7692456 | PEDIATRIC HOSPITAL NO. 2
GIF-H170 | 2726453 | THU DUC CITY HOSPITAL
GIF-LV1(PAL) | 1501505 | LUC NGAN DISTRIC MEDICAL CENTER
GIF-H170 | 2726215 | NATIONAL HOSPITAL OF PEDIATRICS
CF-H170I | 2821841 | UNIVERSITY MEDICAL CENTER HCMC 2
OER-AW | 2431468 | UNIVERSITY MEDICAL CENTER HCMC
GIF-Q150 | 2307377 | QUANG TRI GENERAL HOSPITAL
GIF-Q180 | 2725099 | AN SINH HOSPITAL
CV-190 | 7843005 | HOAN MY CUU LONG HOSPITAL
CLV-190 | 7879682 | HOAN MY CUU LONG HOSPITAL
CF-H170I | 2611044 | TAN PHU HOSPITAL
CV-170 | 7646371 | 30 - 4 HOSPITAL
CF-Q180AI | 2901751 | 30 - 4 HOSPITAL
GIF-HQ190 | 2859761 | 30 - 4 HOSPITAL
GIF-Q180 | 2900986 | 30 - 4 HOSPITAL
GIF-XP190N | 2501414 | TAY NGUYEN REGIONAL GENERAL HOSPITAL
BF-H190 | 2026887 | TAY NGUYEN REGIONAL GENERAL HOSPITAL
GIF-H170 | 2032748 | LAM DONG II GENERAL HOSPITAL
BF-UC180F | 1022402 | CHO RAY HOSPITAL
GIF-Q150 | 2317447 | DA NANG HOSPITAL
GIF-Q150 | 2800627 | VINH MEDICAL UNIVERSITY HOSPITAL
GIF-HQ190 | 2961790 | PHUONG CHAU INTERNATIONAL HOSPITAL
GIF-H170 | 2512643 | VINH MEDICAL UNIVERSITY HOSPITAL
CF-H180AI | 2403077 | BECAMEX INTERNATIONAL HOSPITAL
GIF-H170 | 2032355 | HOAN MY SAIGON HOSPITAL
GIF-EZ1500(E1) | 2203035 | DIGESTIVE & LIVER CENTER
GIF-HQ190 | 2964172 | YERSIN INTERNATIONAL CLINIC
TJF-160VR | 2613432 | DOCTOR TIN CLINIC
MAF-GM2 | 2000210 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
ESG-400 | B002772 | BINH DAN HOSPITAL
GF-UE160-AL5 | 1622727 | BINH DAN HOSPITAL
ENF-P4 | 2653395 | EAR NOSE THROAT (ENT) HOSPITAL HCMC
MAJ-1430 | 7218480 | THU DUC REGIONAL GENERAL HOSPITAL
GIF-XP190N | 2934456 | KIEN GIANG GENERAL HOSPITAL
CLV-190 | 7881254 | YERSIN INTERNATIONAL CLINIC
GIF-Q150 | 2306666 | 331 HOSPITAL
BF-UC180F | 1622337 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-Q150 | 2800888 | AGRICULTURE HOSPITAL
CF-Q180AI | 2202589 | CITY INTERNATIONAL HOSPITAL
CF-H170I | 2821979 | CHO RAY HOSPITAL
CF-H170I | 2822039 | CHO RAY HOSPITAL
ENF-P4 | 2654076 | EAR NOSE THROAT (ENT) HOSPITAL HCMC
ENF-P4 | 2654078 | EAR NOSE THROAT (ENT) HOSPITAL HCMC
CF-Q180AI | 2612827 | VIET - GERMANY AMITY HOSPITAL
CHF-V(STERRAD NX) | 2540905 | VIET - GERMANY AMITY HOSPITAL
CHF-V(STERRAD NX) | 2540907 | VIET - GERMANY AMITY HOSPITAL
GIF-Q180 | 2002295 | VIET - GERMANY AMITY HOSPITAL
CF-H170I | 2822149 | VIET - GERMANY AMITY HOSPITAL
UHI-4 | 7741932 | CITY CHILDREN'S HOSPITAL
GIF-Q150 | 2620602 | THU DUC REGIONAL GENERAL HOSPITAL
CF-Q150I | 2711967 | TRUNG VUONG HOSPITAL
GIF-H170 | 2714740 | THU DUC CITY HOSPITAL
CYF-5A(STERRAD NX) | 2963160 | BINH DAN HOSPITAL
GIF-H170 | 2613041 | QUANG NAM REGIONAL GENERAL HOSPITAL
CF-HQ190I | 2962540 | NGUYEN TRAI HOSPITAL
GIF-H170 | 2726491 | BINH CHANH HOSPITAL
CH-S200-XZ-EB | 7801075 | TIEN GIANG CENTRAL GENERAL HOSPITAL
BF-Q170 | 2900834 | NATIONAL LUNG HOSPITAL
CF-LV1I(PAL) | 2710298 | TAN BINH HOSPITAL
MAJ-1995 | 7804368 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-Q150 | 2307362 | BAO LAM DISTRICT GENERAL HOSPITAL - LAM DONG
CLV-190 | 7099589 | THU DUC REGIONAL GENERAL HOSPITAL
GIF-XP170N | 2931788 | AN SINH HOSPITAL
CYF-5A(STERRAD NX) | 2963044 | BINH DAN HOSPITAL
GIF-Q150 | 2620894 | SAINT PAUL GENERAL HOSPITAL
CF-Q150I | 2611907 | SAINT PAUL GENERAL HOSPITAL
BF-1T150 | 2140443 | PHAM NGOC THACH HOSPITAL
BF-Q170 | 2600259 | E HOSPITAL
CHF-V(STERRAD NX) | 2861837 | THONG NHAT DONG NAI GENERAL HOSPITAL
CF-LV1I(PAL) | 2010688 | LAC VIET FRIENDSHIP HOSPITAL
BF-P180 | 2042172 | PHAM NGOC THACH HOSPITAL
TJF-160VR | 2203177 | BINH AN HOSPITAL
OER-AW | 2732397 | 91 PHAM HUU CHI CLINIC
OER-AW | 2732398 | 91 PHAM HUU CHI CLINIC
A22002A | 641689 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-Q180 | 2304366 | HA TRUNG GENERAL HOSPITAL
CF-H170I | 2610777 | KINH BAC GENERAL HOSPITAL
ENF-P4 | 2038180 | EAR NOSE THROAT (ENT) HOSPITAL HCMC
BF-1T150(STERRAD NX) | 2851878 | UNIVERSITY MEDICAL CENTER HCMC
BF-Q170 | 2800518 | UNIVERSITY MEDICAL CENTER HCMC
GIF-H170 | 2032176 | FAMILY HEALTH CARE JSC
CF-H170I | 2023375 | LE VAN THINH HOSPITAL
GIF-H170 | 2828801 | DIEN BIEN PROVINCIAL GENERAL HOSPITAL
CF-H170I | 2023247 | HOAN MY SAIGON HOSPITAL
TJF-160VR | 2203210 | 108 MILITARY CENTRAL HOSPITAL
CV-170 | 7781942 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CV-190 | 7633849 | DONG NAI 2 HOSPITAL JOINT STOCK COMPANY
GIF-Q150 | 2620934 | COLUMBIA ASIA BINH DUONG INTERNATIONAL HOSPITAL
OTV-S190 | 7704619 | PEDIATRIC HOSPITAL NO. 2
CF-Q150I | 2711924 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-H170 | 2613375 | KINH BAC GENERAL HOSPITAL
URF-V(STERRAD NX) | 2973526 | BINH DAN HOSPITAL
CF-H170I | 2822210 | MILITARY HOSPITAL NO.103
CF-Q150I | 2600012 | MILITARY HOSPITAL NO.103
GIF-H170 | 2032201 | MILITARY HOSPITAL NO.103
CF-Q150I | 2100735 | ONCOLOGY HOSPITAL
CLV-190 | 7987314 | YERSIN INTERNATIONAL CLINIC
CF-H170I | 2611063 | VICTORIA CLINIC
CF-H170I | 2123806 | THONG NHAT HOSPITAL
GIF-H170 | 2930350 | DA NANG HOSPITAL
CF-H170I | 2922671 | YEN BINH GENERAL HOSPITAL(BRANCH OF THAI NGUYEN INTERNATIONAL HOSPITAL)
GIF-HQ190 | 2860244 | KINH BAC GENERAL HOSPITAL
GIF-H170 | 2501688 | 175 MILITARY HOSPITAL
CHF-V2 | 2100217 | UNIVERSITY MEDICAL CENTER HCMC
TJF-160VR | 2503362 | GIA DINH PEOPLE HOSPITAL
CF-Q150I | 2900286 | THONG NHAT DONG NAI GENERAL HOSPITAL
CF-H170I | 2500357 | TIEN GIANG CENTRAL GENERAL HOSPITAL
ESG-400 | B003793 | TIEN GIANG CENTRAL GENERAL HOSPITAL
USG-410 | 7211068 | PHUC HUNG GENERAL HOSPITAL
CF-HQ190I | 2062737 | YERSIN INTERNATIONAL CLINIC
CV-170 | 7543682 | 175 MILITARY HOSPITAL
GIF-Q180 | 2304306 | DONG NAI GENERAL HOSPITAL
ENF-V2 | 2930333 | PEDIATRIC HOSPITAL NO. 2
GIF-H170 | 2032191 | MILITARY HOSPITAL NO.103
GIF-H190 | 2959657 | FV HOSPITAL
CHF-V | 2040123 | GIA DINH PEOPLE HOSPITAL
CF-Q150I | 2401585 | GIA DINH PEOPLE HOSPITAL
TJF-Q180V | 2822930 | DA NANG HOSPITAL
GIF-H170 | 2827514 | HAI PHONG INTERNATIONAL HOSPITAL
GF-UCT180 | 1621409 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-H170 | 2614447 | CHIEM HOA GENERAL HOSPITAL
ESG-400 | 13023W130003 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CLV-190 | 7672663 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
OTV-S7H-1N | 7301597 | SAIGON ENT HOSPITAL
CF-Q150I | 2711965 | ONCOLOGY HOSPITAL
GIF-Q150 | 2721426 | ONCOLOGY HOSPITAL
GIF-Q150 | 2317539 | NINH HOA REGIONAL GENERAL HOSPITAL
GIF-HQ190 | 2963546 | BAI CHAY HOSPITAL
CF-HQ190I | 2761986 | DAI PHUOC GENERAL CLINIC
GIF-Q150 | 2620529 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-XP170N | 2132041 | AN SINH HOSPITAL
GIF-H170 | 2714757 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
BF-1TQ170 | 2800507 | PHAM NGOC THACH HOSPITAL
CF-Q150I | 2501689 | THAI NGUYEN CENTRE GENERAL HOSPITAL
GIF-HQ190 | 2065849 | CA MAU GENERAL HOSPITAL
ENF-P4 | 2936693 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-H170 | 2134590 | HANH PHUC HOSPITAL JOINT STOCK COMPANY
GIF-H170 | 2726447 | THU DUC CITY HOSPITAL
GIF-Q180 | 2001939 | 175 MILITARY HOSPITAL
GIF-H190 | 2856689 | YERSIN INTERNATIONAL CLINIC
CF-Q150I | 2600009 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
BF-1TQ170 | 2400064 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
WM-NP3 | 22112969 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
BF-Q170 | 2000969 | BACH MAI HOSPITAL
CF-H170I | 2721631 | VINH CITY GENERAL HOSPITAL
BF-Q170 | 2700344 | PHAM NGOC THACH HOSPITAL
GIF-H170 | 2033328 | TIEN GIANG CENTRAL GENERAL HOSPITAL
GIF-Q180 | 2102431 | NATIONAL HOSPITAL OF PEDIATRICS
CYF-VA2(STERRAD NX) | 2340649 | SOC TRANG PROVINCE GENERAL HOSPITAL
GIF-EZ1500(E1) | 2102258 | 91 PHAM HUU CHI CLINIC
GIF-H170 | 2726507 | BINH DAN HOSPITAL
GIF-H170 | 2726508 | BINH DAN HOSPITAL
GIF-Q150 | 2721465 | BINH DAN HOSPITAL
CF-H170I | 2023094 | BINH DAN HOSPITAL
CF-HQ190I | 2062806 | BINH DAN HOSPITAL
TJF-160VR | 2400032 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CF-H190I | 2742314 | TAY NGUYEN REGIONAL GENERAL HOSPITAL
GIF-XP170N | 2132030 | BACH MAI HOSPITAL
TJF-150 | 2501737 | NINH THUAN PROVINCE GENERAL HOSPITAL
CF-Q180AI | 2302625 | DONG NAI GENERAL HOSPITAL
CF-Q180AI | 2202503 | DONG NAI GENERAL HOSPITAL
CLV-190 | 7776605 | Y VIET JOINT STOCK COMPANY - UNIVERSITY MEDICAL HOSPITAL CLINIC 1
GIF-Q180 | 2102817 | CAM PHA HOSPITAL
CF-H170I | 2022999 | GIA DINH PEOPLE HOSPITAL
CHF-V2 | 2900120 | THU DUC CITY HOSPITAL
GIF-XP170N | 2931842 | DYM Medical Center Vietnam
ENF-V2(STERRAD NX) | 2660949 | HUE CENTRAL HOSPITAL
OTV-S7PROH-HD-12E | 7130862 | SAIGON ENT HOSPITAL
GIF-HQ190 | 2624018 | 91 PHAM HUU CHI CLINIC
OER-AW | 2220886 | CITY INTERNATIONAL HOSPITAL
CF-H170I | 2610898 | P & T GENERAL HOSPITAL NO.1
CF-H170I | 2922877 | P & T GENERAL HOSPITAL NO.1
CF-H170I | 2721575 | P & T GENERAL HOSPITAL NO.1
CF-H170I | 2023372 | LE VAN THINH HOSPITAL
GIF-HQ190 | 2745220 | VINMEC DA NANG INTERNATIONAL GENERAL HOSPITAL
GIF-XP170N | 2500720 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CF-H170I | 2922882 | P & T GENERAL HOSPITAL NO.1
CF-H170I | 2610979 | SON TAY DISTRICT GENERAL HOSPITAL
CF-HQ190I | 2962541 | DIGESTIVE & LIVER CENTER
BF-1TQ170 | 2101424 | PHAM NGOC THACH HOSPITAL
ESG-400 | 16074W140008 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
BF-1TQ170 | 2700372 | PHAM NGOC THACH HOSPITAL
GIF-H170 | 2827522 | HAI PHONG INTERNATIONAL HOSPITAL
CF-H170I | 2922686 | HAI PHONG INTERNATIONAL HOSPITAL
GIF-Q150 | 2721036 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
ENF-P4 | 2038182 | EAR NOSE THROAT (ENT) HOSPITAL HCMC
GIF-H170 | 2613373 | BAC NINH PROVINCAL GENERAL HOSPITAL
GIF-H170 | 2032776 | HA TINH GENERAL HOSPITAL
CF-Q150I | 2501717 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-H170 | 2512425 | PHAM NGOC THACH UNIVERSITY OF MEDICINE
GIF-H170 | 2726509 | BINH DAN HOSPITAL
GIF-H170 | 2726512 | BINH DAN HOSPITAL
GIF-H170 | 2031766 | BINH DAN HOSPITAL
CF-H170I | 2023093 | BINH DAN HOSPITAL
CF-H170I | 2023095 | BINH DAN HOSPITAL
WA2T430A | 817832 | HUE CENTRAL HOSPITAL - BRANCH 2
GIF-H170 | 2826962 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
GIF-HQ190 | 2963034 | BAI CHAY HOSPITAL
CF-H170I | 2922919 | E HOSPITAL
GIF-H190 | 2959684 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CF-H170I | 2300120 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
GIF-H170 | 2501726 | 15 MILITARY HOSPITAL
CF-H170I | 2922988 | HOAN MY SAIGON HOSPITAL
GIF-Q150 | 2002397 | GO CONG REGIONAL GENERAL HOSPITAL
CV-170 | 7647092 | GO CONG REGIONAL GENERAL HOSPITAL
MAJ-1430 | 7013003 | GO CONG REGIONAL GENERAL HOSPITAL
GIF-HQ190 | 2168841 | QUANG NINH GENERAL HOSPITAL
TJF-160VR | 2403330 | CAN THO CITY GENERAL HOSPITAL
CF-H170I | 2610812 | 108 MILITARY CENTRAL HOSPITAL
GIF-H180 | 2900797 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
TJF-160VR | 2303240 | DONG NAI GENERAL HOSPITAL
GIF-Q150 | 2922256 | GIA DINH PEOPLE HOSPITAL
GIF-H170 | 2828765 | VIET NAM-THUY DIEN UONG BI HOSPITAL
CF-H190I | 2842889 | HA NOI FRENCH HOSPITAL
GIF-XP170N | 2031954 | BINH TAN HOSPITAL
TJF-Q180V | 2923466 | CAN THO CITY GENERAL HOSPITAL
GIF-EZ1500(E1) | 2203060 | 91 PHAM HUU CHI CLINIC
GIF-H180 | 2307461 | CITY INTERNATIONAL HOSPITAL
GIF-H170 | 2613645 | THUAN THANH DISTRICT HOSPITAL
GIF-Q180 | 2314635 | THONG NHAT HOSPITAL
CF-Q150I | 2000470 | HUNG YEN GENERAL HOSPITAL
GIF-H180 | 2106658 | COLUMBIA ASIA BINH DUONG INTERNATIONAL HOSPITAL
GIF-H170 | 2827545 | AMERICAN INTERNATIONAL HOSPITAL
GIF-Q150 | 2821655 | UNIVERSITY MEDICAL CENTER HCMC 2
GIF-H170 | 2726291 | P & T GENERAL HOSPITAL NO.1
GIF-H190 | 2960244 | YERSIN INTERNATIONAL CLINIC
BF-1T150 | 2600031 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CF-H170I | 2721442 | HONG NGOC GENERAL HOSPITAL
GIF-Q150 | 2800556 | HONG DUC GENERAL HOSPITAL
CF-H170I | 2611049 | CU CHI DISTRICT HOSPITAL
A70962A | 807638 | GIA DINH GENERAL HOSPITAL
GIF-Q150 | 2104639 | UNIVERSITY MEDICAL CENTER HCMC
CLV-190 | 7665844 | UNIVERSITY MEDICAL CENTER HCMC
CLV-190 | 7983341 | UNIVERSITY MEDICAL CENTER HCMC
GIF-H170 | 2401534 | HONG NGOC GENERAL HOSPITAL
GIF-HQ190 | 2064724 | HONG NGOC GENERAL HOSPITAL
TJF-Q180V | 2822247 | BAC LIEU PROVINCE GENERAL HOSPITAL
CHF-V(STERRAD NX) | 2861701 | 108 MILITARY CENTRAL HOSPITAL
CF-H170I | 2610694 | QUANG NAM REGIONAL GENERAL HOSPITAL
CV-180 | 7008914 | 175 MILITARY HOSPITAL
WA2T430A | 754564 | UNIVERSITY MEDICAL CENTER HCMC
WA29042A | 800364 | UNIVERSITY MEDICAL CENTER HCMC
GIF-Q180 | 2102843 | CAM PHA HOSPITAL
ESG-400 | B007487 | TIEN GIANG CENTRAL GENERAL HOSPITAL
GIF-Q150 | 2901178 | ONCOLOGY HOSPITAL
CF-HQ190I | 2062752 | VINMEC CENTRAL PARK HOSPITAL
ENF-P4 | 2256573 | EAR NOSE THROAT (ENT) HOSPITAL HCMC
OTV-SC2 | 7911348 | HOAN MY DA NANG GENERAL HOSPITAL
GIF-H170 | 2929430 | MEDLATEC GENERAL HOSPITAL
WA2T430A | 811235 | UNIVERSITY MEDICAL CENTER HCMC
TJF-150 | 2100695 | 108 MILITARY CENTRAL HOSPITAL
TJF-Q180V | 2721549 | 108 MILITARY CENTRAL HOSPITAL
CF-Q150I | 2201064 | UNIVERSITY MEDICAL CENTER HCMC
GIF-Q150 | 2306605 | UNIVERSITY MEDICAL CENTER HCMC
GIF-H170 | 2501640 | UNIVERSITY MEDICAL CENTER HCMC
GIF-Q150 | 2306547 | UNIVERSITY MEDICAL CENTER HCMC
CF-Q180AI | 2901620 | HA TINH GENERAL HOSPITAL
CF-Q150I | 2200786 | TAY NGUYEN REGIONAL GENERAL HOSPITAL
ESG-100 | 12332P500001 | HONG DUC GENERAL HOSPITAL
TJF-Q180V | 2135083 | CAN THO UNIVERSITY OF MEDICINE AND PHARMACY HOSPITAL
ESG-400 | B006224 | NGUYEN DINH CHIEU HOSPITAL
CH-S200-XZ-EB | 7912012 | NGUYEN DINH CHIEU HOSPITAL
GIF-H170 | 2930005 | BINH THANH HOSPITAL
TJF-Q180V | 2618914 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CF-Q150I | 2201066 | UNIVERSITY MEDICAL CENTER HCMC
CF-Q150I | 2201068 | UNIVERSITY MEDICAL CENTER HCMC
CF-H170I | 2610672 | UNIVERSITY MEDICAL CENTER HCMC
CF-H170I | 2500359 | UNIVERSITY MEDICAL CENTER HCMC
GIF-H170 | 2828800 | MILITARY HOSPITAL NO.103
CHF-V(STERRAD NX) | 2540808 | CAN THO CITY GENERAL HOSPITAL
TJF-Q180V | 2934260 | BINH DAN HOSPITAL
CF-HQ190I | 2163024 | UNIVERSITY MEDICAL CENTER HCMC
GIF-HQ190 | 2858918 | UNIVERSITY MEDICAL CENTER HCMC
BF-1TQ170 | 2800579 | CHO RAY HOSPITAL
CF-HQ190I | 2962645 | UNIVERSITY MEDICAL CENTER HCMC
CF-HQ190I | 2862177 | UNIVERSITY MEDICAL CENTER HCMC
ENF-VH2 | 2100596 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CF-H170I | 2822254 | 175 MILITARY HOSPITAL
GIF-H190 | 2856700 | YERSIN INTERNATIONAL CLINIC
GIF-XP190N | 2934446 | KIEN GIANG GENERAL HOSPITAL
OTV-S400 | 7500360 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-H170 | 2032182 | DISTRICT 1 HOSPITAL
CF-HQ190I | 2651604 | FV HOSPITAL
BF-XP190 | 2500336 | HAI PHONG CHILDREN'S HOSPITAL
ENF-VH2 | 2000558 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-H170 | 2614313 | SHINGMARK MEDICAL UNIVERSITY HOSPITAL
CLV-190 | 7457730 | UNIVERSITY MEDICAL CENTER HCMC
CLV-190 | 7878415 | UNIVERSITY MEDICAL CENTER HCMC
GIF-Q150 | 2821551 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
GIF-H170 | 2614708 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
GIF-Q150 | 2620429 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
WM-NP3 | 22217464 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-H170 | 2828766 | VIET NAM-THUY DIEN UONG BI HOSPITAL
GIF-H170 | 2828749 | VIET NAM-THUY DIEN UONG BI HOSPITAL
GIF-H170 | 2828428 | VICTORIA CLINIC
CF-Q150I | 2000304 | PUBLIC SECURITY MINISTRY 19/8 HOSPITAL
GIF-Q180 | 2900977 | PUBLIC SECURITY MINISTRY 19/8 HOSPITAL
CF-Q180AI | 2901750 | PUBLIC SECURITY MINISTRY 19/8 HOSPITAL
CF-Q150I | 2501826 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
TJF-160VR | 2400035 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-Q150 | 2418084 | TAM BINH DISTRICT GENERAL HOSPITAL
CV-180 | 7206151 | TAM BINH DISTRICT GENERAL HOSPITAL
URF-P5(STERRAD NX) | 2856781 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
TJF-Q180V | 2102168 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CF-HQ190I | 2962539 | NGUYEN TRAI HOSPITAL
CF-H170I | 2123829 | UNIVERSITY MEDICAL CENTER HCMC 2
CF-H170I | 2821933 | SAI GON NHA TRANG GENERAL HOSPITAL
CF-Q150I | 2201031 | THU DUC REGIONAL GENERAL HOSPITAL
GIF-H170 | 2032186 | DISTRICT 1 HOSPITAL
GIF-Q150 | 2922251 | GIA DINH PEOPLE HOSPITAL
GIF-Q150 | 2519534 | GIA DINH PEOPLE HOSPITAL
CLV-S400 | 7600670 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
UHI-4 | 7649001 | GO CONG REGIONAL GENERAL HOSPITAL
WM-NP2 | 21736722 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CV-170 | 7432694 | SAIGON ENT HOSPITAL
BF-1T150(STERRAD NX) | 2751775 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CF-H170I | 2711404 | ANH KHOA MEDICAL EQUIPMENT Co., LTD
CV-170 | 7779918 | ANH KHOA MEDICAL EQUIPMENT Co., LTD
CF-HQ190L | 2185916 | VINH LONG GENERAL HOSPITAL
CF-EZ1500DI(E1) | 2101288 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CF-HQ190I | 2200170 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CF-H170I | 2610885 | BAC GIANG ONCOLOGY HOSPITAL
CLV-190 | 7667392 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-Q150 | 2821548 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
TJF-Q180V | 2034461 | CHO RAY HOSPITAL
CF-HQ190L | 2513177 | FAMILY MEDICAL PRACTICE HCMC CLINIC
CHF-P60 | 2141146 | UNIVERSITY MEDICAL CENTER HCMC
CV-170 | 7884739 | 175 MILITARY HOSPITAL
TJF-150 | 2501754 | DA NANG C HOSPITAL
CF-Q150L | 2001283 | GO CONG REGIONAL GENERAL HOSPITAL
CF-H190I | 2621656 | SAINT PAUL GENERAL HOSPITAL
CF-H190I | 2621704 | SAINT PAUL GENERAL HOSPITAL
CF-H180AI | 2901379 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CF-H170I | 2821872 | VICTORIA CLINIC
GIF-H170 | 2614706 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
ENF-V2(STERRAD NX) | 2450762 | PEDIATRIC HOSPITAL NO. 2
TJF-Q180V | 2822572 | 30 - 4 HOSPITAL
CF-H170I | 2500401 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-H170 | 2929187 | GIA AN 115 HOSPITAL
CF-H170I | 2721510 | MEDLATEC GENERAL HOSPITAL
CYF-5A(STERRAD NX) | 2963159 | BINH DAN HOSPITAL
WA2T430A | 800276 | BINH DAN HOSPITAL
WA2T430A | 801878 | BINH DAN HOSPITAL
CH-S190-08-LB | 7705506 | BINH DAN HOSPITAL
GIF-H170 | 2031763 | BINH DAN HOSPITAL
GIF-H170 | 2031767 | BINH DAN HOSPITAL
GIF-H170 | 2031770 | BINH DAN HOSPITAL
GIF-Q150 | 2519871 | CHO RAY HOSPITAL
GIF-Q150 | 2721222 | CHO RAY HOSPITAL
GIF-H170 | 2512428 | CHO RAY HOSPITAL
CF-H170I | 2500462 | CHO RAY HOSPITAL
GIF-H170 | 2827838 | CHO RAY HOSPITAL
GIF-H170 | 2827847 | CHO RAY HOSPITAL
GIF-H170 | 2828976 | CHO RAY HOSPITAL
GIF-H170 | 2828962 | CHO RAY HOSPITAL
CF-H180AI | 2102516 | COLUMBIA ASIA BINH DUONG INTERNATIONAL HOSPITAL
TJF-160VR | 2500233 | UNIVERSITY MEDICAL CENTER HCMC
GIF-1TH190 | 2942855 | FV HOSPITAL
TJF-Q180V | 2822714 | FV HOSPITAL
CF-EZ1500DI(E1) | 2101282 | TAM ANH HOSPITAL- HCMC BRANCH
CF-H190I | 2842982 | AN GIANG REGIONAL GENERAL HOSPITAL
CF-Q150I | 2401573 | NATIONAL TRANSPORT HOSPITAL
GIF-H170 | 2512452 | HAI DUONG UNIVERSITY OF MEDICAL TECHNOLOGY
CV-170 | 7010211 | THIEN PHU CLINIC
CF-H170I | 2500461 | BINH DINH PROVINCE GENERAL HOSPITAL
CF-H170I | 2822152 | VICTORIA CLINIC
GIF-XP180N | 2501683 | AN SINH HOSPITAL
CF-H170I | 2922878 | 175 MILITARY HOSPITAL
CF-H170I | 2922880 | 175 MILITARY HOSPITAL
OTV-S7PROH-HD-12E | 7646020 | LE VAN THINH HOSPITAL
CLV-180 | 7315511 | HUONG TRA MEDICAL CENTER
GIF-Q150 | 2307376 | HUONG TRA MEDICAL CENTER
OFP-2 | 21863256 | FV HOSPITAL
CF-HQ190I | 2651622 | 91 PHAM HUU CHI CLINIC
GIF-H170 | 2135352 | PUBLIC SECURITY MINISTRY 19/8 HOSPITAL
CLV-190 | 7103415 | UNIVERSITY MEDICAL CENTER HCMC
GIF-Q150 | 2600028 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-HQ190 | 2270495 | HOAN MY BINH PHUOC HOSPITAL
CHF-V | 2651166 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CV-V1 | 1200268 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-Q150 | 2104661 | UNIVERSITY MEDICAL CENTER HCMC
CF-HQ190I | 2962505 | UNIVERSITY MEDICAL CENTER HCMC
GIF-HQ190 | 2858921 | UNIVERSITY MEDICAL CENTER HCMC
GIF-Q150 | 2204665 | UNIVERSITY MEDICAL CENTER HCMC
CF-HQ190I | 2862182 | UNIVERSITY MEDICAL CENTER HCMC
CF-HQ190I | 2962646 | UNIVERSITY MEDICAL CENTER HCMC
UES-40 | 7359598 | CITY INTERNATIONAL HOSPITAL
LF-DP | 2024653 | PHAM NGOC THACH HOSPITAL
ENF-P4 | 2256573 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CV-190 | 7841405 | 175 MILITARY HOSPITAL
CLV-190 | 7878420 | 175 MILITARY HOSPITAL
CF-H170I | 2500365 | 175 MILITARY HOSPITAL
GIF-HQ190 | 2964226 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-HQ190 | 2510896 | FAMILY MEDICAL PRACTICE HCMC CLINIC
GIF-H170 | 2032350 | HOAN MY SAIGON HOSPITAL
GIF-H170 | 2726297 | MEDLATEC GENERAL HOSPITAL
CF-H190I | 2511308 | VIET - GERMANY AMITY HOSPITAL
UHI-4 | 7638715 | THU DUC CITY HOSPITAL
GIF-Q150 | 2418200 | THANH NHAN HOSPITAL
ENF-P4 | 2038184 | EAR NOSE THROAT (ENT) HOSPITAL HCMC
ENF-P4 | 2038183 | EAR NOSE THROAT (ENT) HOSPITAL HCMC
ENF-P4 | 1913035 | EAR NOSE THROAT (ENT) HOSPITAL HCMC
BF-XP190 | 2012029 | NATIONAL HOSPITAL OF PEDIATRICS
CF-H190I | 2621703 | SAINT PAUL GENERAL HOSPITAL
GIF-Q150 | 2821673 | SAINT PAUL GENERAL HOSPITAL
GIF-H170 | 2614348 | VICTORIA CLINIC
CF-HQ190I | 2862146 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-H170 | 2133749 | THAI NGUYEN CENTRE GENERAL HOSPITAL
BF-H190 | 2026703 | MILITARY HOSPITAL NO.103
GIF-H190 | 2629301 | SAINT PAUL GENERAL HOSPITAL
GIF-H190 | 2629187 | SAINT PAUL GENERAL HOSPITAL
GIF-Q150 | 2721421 | GIA DINH PEOPLE HOSPITAL
GIF-Q180 | 2900553 | CAO BANG GENERAL HOSPITAL
GIF-H170 | 2929099 | DISTRICT 4 HOSPITAL
GIF-H170 | 2931031 | 175 MILITARY HOSPITAL
BF-1TQ170 | 2800630 | 175 MILITARY HOSPITAL
BF-1TH190 | 2227631 | TAM ANH HOSPITAL- HCMC BRANCH
GIF-LV1(PAL) | 2811598 | LE VAN VIET HOSPITAL
GIF-H170 | 2134671 | HOP LUC GENERAL HOSPITAL
OTV-S7PROH-HD-10E | 7800454 | UNIVERSITY MEDICAL CENTER HCMC
USG-400 | 9507582 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
ENF-V3 | 2612331 | PEDIATRIC HOSPITAL NO. 2
BF-P190 | 2601165 | PEDIATRIC HOSPITAL NO. 2
UHI-4 | 7649125 | GIA DINH PEOPLE HOSPITAL
GIF-Q150 | 2419031 | NATIONAL HOSPITAL OF PEDIATRICS
GIF-H170 | 2512765 | THAI BINH MEDICAL UNIVERSITY HOSPITAL
CF-H190I | 2043565 | THONG NHAT HOSPITAL
CF-H170I | 2822255 | DIEN BIEN PROVINCIAL GENERAL HOSPITAL
GIF-EZ1500(E1) | 2203400 | YERSIN INTERNATIONAL CLINIC
GIF-HQ190 | 2201619 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CF-H170I | 2821916 | AMERICAN INTERNATIONAL HOSPITAL
GIF-HQ190 | 2510394 | E HOSPITAL
GIF-H170 | 2032019 | MILITARY HOSPITAL NO.103
GIF-H170 | 2930318 | YEN BINH GENERAL HOSPITAL(BRANCH OF THAI NGUYEN INTERNATIONAL HOSPITAL)
GIF-H170 | 2032200 | MILITARY HOSPITAL NO.103
CF-HQ190I | 2062831 | HOAN MY SAIGON HOSPITAL
GIF-Q150 | 2721249 | TRUNG VUONG HOSPITAL
OTV-S7H-1N | 7526949 | AN GIANG CENTRAL GENERAL HOSPITAL
OTV-S7H-1N | 7526948 | AN GIANG CENTRAL GENERAL HOSPITAL
BF-1TQ170 | 2700427 | GIA DINH PEOPLE HOSPITAL
CF-H170I | 2300122 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
UHI-3 | 1200128 | EASTERN ARMY AND PEOPLE HOSPITAL
GIF-H170 | 2827811 | CHO RAY HOSPITAL
CF-H170I | 2822044 | CHO RAY HOSPITAL
CF-H170I | 2023173 | MILITARY HOSPITAL NO.103
MAJ-1430 | 7015140 | 175 MILITARY HOSPITAL
CF-H170I | 2611064 | VICTORIA CLINIC
TJF-Q180V | 2923209 | GIA AN 115 HOSPITAL
GIF-Q150 | 2922294 | HAI PHONG INTERNATIONAL HOSPITAL
BF-1T150 | 2040388 | PHAM NGOC THACH HOSPITAL
CF-H190I | 2842983 | YERSIN INTERNATIONAL CLINIC
GIF-EZ1500(E1) | 2203338 | YERSIN INTERNATIONAL CLINIC
CF-Q180AI | 2302638 | E HOSPITAL
GIF-Q150 | 2519893 | E HOSPITAL
GIF-Q150 | 2002973 | E HOSPITAL
GIF-Q150 | 2510024 | E HOSPITAL
CF-HQ190I | 2651647 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-H170 | 2930180 | NAM SAI GON INTERNATIONAL GENERAL HOSPITAL
GIF-H170 | 2726159 | YEN BAI PROVINCIAL GENERAL HOSPITAL
GIF-H170 | 2726158 | YEN BAI PROVINCIAL GENERAL HOSPITAL
CF-H170I | 2610729 | TIEN DU MEDICAL CENTER
GIF-H190 | 2629257 | SAINT PAUL GENERAL HOSPITAL
GIF-H190 | 2629277 | SAINT PAUL GENERAL HOSPITAL
GIF-H190 | 2062408 | HUNG VIET CANCER HOSPITAL
CLV-S190 | 7724888 | ONCOLOGY HOSPITAL
GIF-H170 | 2929174 | CITY INTERNATIONAL HOSPITAL
PCF-Q180AI | 2200613 | CITY INTERNATIONAL HOSPITAL
GIF-H170 | 2512255 | MEDIC SAIGON BENTRE GENERAL CLINIC
CF-H170I | 2023004 | TRIEU AN HOSPITAL
CF-H170I | 2500532 | E HOSPITAL
CHF-V(STERRAD NX) | 2162662 | 108 MILITARY CENTRAL HOSPITAL
GIF-H170 | 2726214 | NATIONAL HOSPITAL OF PEDIATRICS
BF-1T150(STERRAD NX) | 2751830 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-H170 | 2931037 | THANG BINH MEDICAL CENTER
CV-70 | 7120964 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
OTV-S7H-1N | 7200631 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CV-190 | 7401658 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CLV-190 | 7451079 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-H190 | 2062028 | E HOSPITAL
GIF-H190 | 2062055 | E HOSPITAL
BF-1TQ170 | 2600246 | THONG NHAT HOSPITAL
ESG-400 | 15287W140010 | THONG NHAT HOSPITAL
WA2T430A | 805852 | KHANH HOA PROVINCE GENERAL HOSPITAL
WA2T412A | 807204 | KHANH HOA PROVINCE GENERAL HOSPITAL
GIF-HQ190 | 2270403 | HOAN MY BINH PHUOC HOSPITAL
TJF-160VR | 2400129 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-H170 | 2612990 | UNIVERSITY MEDICAL CENTER HCMC
GIF-HQ190 | 2858917 | UNIVERSITY MEDICAL CENTER HCMC
GIF-HQ190 | 2858926 | UNIVERSITY MEDICAL CENTER HCMC
GIF-HQ190 | 2858915 | UNIVERSITY MEDICAL CENTER HCMC
GIF-HQ190 | 2963804 | UNIVERSITY MEDICAL CENTER HCMC
GIF-Q180 | 2304303 | UNIVERSITY MEDICAL CENTER HCMC
GIF-Q150 | 2204663 | UNIVERSITY MEDICAL CENTER HCMC
GIF-HQ190 | 2963796 | UNIVERSITY MEDICAL CENTER HCMC
GIF-Q150 | 2306585 | UNIVERSITY MEDICAL CENTER HCMC
CF-HQ190I | 2862180 | UNIVERSITY MEDICAL CENTER HCMC
CF-HQ190I | 2163153 | UNIVERSITY MEDICAL CENTER HCMC
GIF-HQ190 | 2858922 | UNIVERSITY MEDICAL CENTER HCMC
CF-Q150I | 2201063 | UNIVERSITY MEDICAL CENTER HCMC
OTV-S400 | 7711475 | 175 MILITARY HOSPITAL
BF-P190 | 2912971 | 108 MILITARY CENTRAL HOSPITAL
BF-H190 | 2026869 | 108 MILITARY CENTRAL HOSPITAL
GIF-Q150 | 2700057 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CH-S190-XZ-E | 7707828 | BINH TAN HOSPITAL
BF-1T150(STERRAD NX) | 2341161 | BINH THUAN LUNG HOSPITAL
WA4KL130 | W778920 | HUE CENTRAL HOSPITAL
CF-HQ190I | 2962610 | BAI CHAY HOSPITAL
CHF-V2 | 2200299 | BINH DINH PROVINCE GENERAL HOSPITAL
CH-S400-XZ-EB | 7702554 | 175 MILITARY HOSPITAL
GIF-Q150 | 2721447 | THONG NHAT HOSPITAL
BF-H190 | 2026062 | CHO RAY HOSPITAL
CF-HQ190I | 2862179 | BAI CHAY HOSPITAL
CF-V70L-PAL | 2810920 | ARMY HOSPITAL - DONG THAP
CF-H170I | 2023179 | MILITARY HOSPITAL NO.103
CF-H170I | 2023197 | DISTRICT 1 HOSPITAL
GIF-XP170N | 2831693 | AN SINH HOSPITAL
ENF-V4 | 2200643 | MEDLATEC GENERAL HOSPITAL
GIF-XP150N | 2512519 | HAI CHAU GENERAL HOSPITAL
CF-H170I | 2922606 | NAM SAI GON INTERNATIONAL GENERAL HOSPITAL
GIF-HQ190 | 2963822 | UNIVERSITY MEDICAL CENTER HCMC
GIF-Q150 | 2205895 | UNIVERSITY MEDICAL CENTER HCMC
GIF-EZ1500(E1) | 2203322 | YERSIN INTERNATIONAL CLINIC
GIF-Q150 | 2721449 | GIA DINH PEOPLE HOSPITAL
GIF-H170 | 2032024 | MILITARY HOSPITAL NO.103
CF-H170I | 2023178 | MILITARY HOSPITAL NO.103
GIF-V70-PAL | 2512037 | ARMY HOSPITAL - DONG THAP
ESG-400 | B008141 | TIEN GIANG CENTRAL GENERAL HOSPITAL
CF-H170I | 2821772 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-Q150 | 2002863 | CAN THO CARDIOLOGY HOSPITAL
GIF-Q150 | 2204943 | CAN THO CARDIOLOGY HOSPITAL
GIF-H170 | 2929030 | THAO NGUYEN GENERAL HOSPITAL
BF-1T150(STERRAD NX) | 2441384 | 108 MILITARY CENTRAL HOSPITAL
UHI-4 | 7322342 | 175 MILITARY HOSPITAL
CF-Q180AI | 2002022 | 175 MILITARY HOSPITAL
GIF-Q180 | 2001933 | 175 MILITARY HOSPITAL
BF-XP190 | 2011950 | NATIONAL HOSPITAL OF PEDIATRICS
CF-H170I | 2023234 | DOCTOR CHUNG CLINIC
CF-H170I | 2611087 | LE VAN THINH HOSPITAL
GIF-H190 | 2959618 | FV HOSPITAL
CF-H170I | 2822214 | VIET NAM-THUY DIEN UONG BI HOSPITAL
GIF-HQ190 | 2756960 | DA NANG HOSPITAL
CYF-5A(STERRAD NX) | 2263329 | BINH DAN HOSPITAL
GIF-HQ190 | 2373011 | NGUYEN TRAI HOSPITAL
BF-Q170 | 2700416 | 108 MILITARY CENTRAL HOSPITAL
GIF-H170 | 2930558 | 108 MILITARY CENTRAL HOSPITAL
GIF-H170 | 2930550 | 108 MILITARY CENTRAL HOSPITAL
CF-Q150I | 2200978 | PHUC THO GENERAL HOSPITAL
CHF-V(STERRAD NX) | 2962172 | 175 MILITARY HOSPITAL
WM-NP3 | 22217462 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CF-EZ1500DI(E1) | 2101255 | 91 PHAM HUU CHI CLINIC
MAJ-1430 | 7831025 | UNIVERSITY MEDICAL CENTER HCMC 2
GIF-HQ190 | 2961587 | UNIVERSITY MEDICAL CENTER HCMC
CF-H170I | 2922523 | THU DUC REGIONAL GENERAL HOSPITAL
GIF-H170 | 2032249 | DOCTOR CHUNG CLINIC
CF-H170I | 2822213 | VIET NAM-THUY DIEN UONG BI HOSPITAL
CF-H170I | 2500553 | 108 MILITARY CENTRAL HOSPITAL
TJF-160VR | 2513421 | THONG NHAT DONG NAI GENERAL HOSPITAL
OTV-S7PROH-HD-12E | 7645977 | XUYEN A GENERAL HOSPITAL HCMC
CLV-180 | 7113176 | UNIVERSITY MEDICAL CENTER HCMC
ESG-400 | B006290 | HUNG VUONG HOSPITAL
GIF-H170 | 2931454 | HOAN MY SAIGON HOSPITAL
GIF-H170 | 2512767 | 108 MILITARY CENTRAL HOSPITAL
GIF-H170 | 2613504 | 108 MILITARY CENTRAL HOSPITAL
GIF-H170 | 2613513 | 108 MILITARY CENTRAL HOSPITAL
GIF-H170 | 2726045 | 108 MILITARY CENTRAL HOSPITAL
CF-Q150I | 2501701 | E HOSPITAL
CF-HQ190I | 2761949 | FV HOSPITAL
GIF-Q150 | 2600011 | LOTUS CLINIC
ENF-P4 | 2653398 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
OTV-S7H-1N | 7717585 | HOAN MY CUU LONG HOSPITAL
GIF-H170 | 2133612 | NAM SAI GON INTERNATIONAL GENERAL HOSPITAL
CF-Q150I | 2601853 | 108 MILITARY CENTRAL HOSPITAL
CF-H170I | 2500486 | THANH NHAN HOSPITAL
CF-H170I | 2821770 | TRAN-LE TRADING AND PRODUCTION CO., LTD
GIF-Q180 | 2001483 | BAI CHAY HOSPITAL
GIF-HQ190 | 2860311 | YERSIN INTERNATIONAL CLINIC
GIF-Q150 | 2700116 | THONG NHAT DONG NAI GENERAL HOSPITAL
OTV-S7PRO | 7579135 | FV HOSPITAL
WM-NP2 | 21404608 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
WM-NP2 | 21605682 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CF-HQ190I | 2200172 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CF-H170I | 2123720 | THAI NGUYEN INTERNATIONAL HOSPITAL
ENF-V4 | 2100555 | NAM SAI GON INTERNATIONAL GENERAL HOSPITAL
CF-Q150I | 2600011 | 175 MILITARY HOSPITAL
CF-HQ190I | 2761954 | 175 MILITARY HOSPITAL
GIF-H170 | 2613509 | 108 MILITARY CENTRAL HOSPITAL
BF-P150 | 2751187 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CF-H170I | 2822236 | CHO RAY HOSPITAL
GIF-H170 | 2827828 | CHO RAY HOSPITAL
GIF-H170 | 2828966 | CHO RAY HOSPITAL
CV-190 | 7504785 | CHO RAY HOSPITAL
CF-H170I | 2023007 | HOAN MY THU DUC INTERNATIONAL GENERAL HOSPITAL
A70960A | 745490 | DIAMOND HIGH QUALITY MEDICINE
CF-H180AI | 2901375 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
GIF-Q150 | 2519523 | HUNG YEN GENERAL HOSPITAL
GIF-H170 | 2614449 | LE VAN THINH HOSPITAL
CF-H190I | 2842922 | HA NOI FRENCH HOSPITAL
GIF-Q150 | 2418984 | GIA DINH PEOPLE HOSPITAL
OER-AW | 2732344 | Y VIET JOINT STOCK COMPANY - UNIVERSITY MEDICAL HOSPITAL CLINIC 1
GIF-H170 | 2931356 | HOAN MY SAIGON HOSPITAL
GIF-H180 | 2307476 | BAI CHAY HOSPITAL
CF-Q150I | 2711923 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CF-EZ1500DI(E1) | 2000687 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-LV1(PAL) | 1300547 | VINH PHUC TRANSPORTATION HOSPITAL
CF-H170I | 2610732 | YEN PHONG DISTRICT HOSPITAL
CF-EZ1500DI(E1) | 2000475 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
UCR | 7734897 | VICTORIA CLINIC
GIF-H170 | 2032460 | LE VAN THINH HOSPITAL
ENF-P4 | 2855200 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-Q150 | 2821550 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-H170 | 2613141 | QUE VO GENERAL HOSPITAL
LTF-S190-5 | 2723727 | CITY CHILDREN'S HOSPITAL
CF-H170I | 2922728 | 108 MILITARY CENTRAL HOSPITAL
GIF-Q150 | 2721266 | 108 MILITARY CENTRAL HOSPITAL
GIF-HQ190 | 2612435 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
PCF-Q180AI | 2300634 | QUANG NINH OBSTETRICS AND PEDIATRICS HOSPITAL
GIF-H170 | 2032189 | MILITARY HOSPITAL NO.103
CF-Q150I | 2501751 | CHO RAY HOSPITAL
CF-H170I | 2500459 | CHO RAY HOSPITAL
GIF-H170 | 2032267 | LE VAN THINH HOSPITAL
BF-P190 | 2711363 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
TJF-Q180V | 2934121 | 175 MILITARY HOSPITAL
USG-400 | 9710689 | TAN PHU HOSPITAL
CF-H170I | 2821759 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
ENF-V4 | 2100558 | NAM SAI GON INTERNATIONAL GENERAL HOSPITAL
CF-H170I | 2922730 | 108 MILITARY CENTRAL HOSPITAL
GIF-Q150 | 2620618 | 108 MILITARY CENTRAL HOSPITAL
GIF-H170 | 2726044 | 108 MILITARY CENTRAL HOSPITAL
CLV-S40 | 7017042 | CAN THO PEDIATRIC HOSPITAL
CF-H170I | 2821871 | VICTORIA CLINIC
GIF-H170 | 2828967 | CHO RAY HOSPITAL
CF-HQ190I | 2631375 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CLV-S190 | 7202684 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-H170 | 2726047 | 108 MILITARY CENTRAL HOSPITAL
GIF-H170 | 2827008 | TRAN-LE TRADING AND PRODUCTION CO., LTD
CF-H190I | 2943385 | PHUONG DONG GENERAL HOSPITAL
CF-Q150I | 2601852 | 108 MILITARY CENTRAL HOSPITAL
GIF-Q150 | 2800837 | 108 MILITARY CENTRAL HOSPITAL
CF-H170I | 2123731 | TIEN GIANG CENTRAL GENERAL HOSPITAL
GIF-H170 | 2033336 | TIEN GIANG CENTRAL GENERAL HOSPITAL
TJF-Q180V | 2822502 | AMERICAN INTERNATIONAL HOSPITAL
CLV-190 | 7000611 | TAN PHU HOSPITAL
GIF-Q150 | 2721034 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CV-190 | 7271882 | HOAN MY BINH PHUOC HOSPITAL
GIF-H190 | 2855924 | PHUONG DONG GENERAL HOSPITAL
CF-HQ190I | 2631330 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CF-H170I | 2922458 | MEDLATEC GENERAL HOSPITAL
BF-Q170 | 2101659 | FAMILY HEALTH CARE JSC
BF-Q170 | 2001072 | PHAM NGOC THACH HOSPITAL
GIF-H190 | 2164711 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
CV-150 | 7831741 | CAN THO CARDIOLOGY HOSPITAL
MAJ-1558 | 7802048 | CAN THO CARDIOLOGY HOSPITAL
GIF-H170 | 2614265 | TAN PHU HOSPITAL
BF-1TQ170 | 2800499 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CV-190 | 7211779 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CF-H170I | 2610671 | MEDLATEC GENERAL HOSPITAL
ENF-VH2 | 2200665 | VINMEC CENTRAL PARK HOSPITAL
CF-Q150I | 2301342 | 108 MILITARY CENTRAL HOSPITAL
EU-ME2 | 7501333 | DIGESTIVE & LIVER CENTER
GIF-H170 | 2338192 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CF-H170I | 2610981 | TAN CANG CLINIC
ESG-150 | B000785 | DOCTOR CHUNG CLINIC
BF-1T150(STERRAD NX) | 2441380 | BACH MAI HOSPITAL
GIF-HQ190 | 2201606 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
GIF-H180 | 2900799 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CF-H170I | 2611115 | LAM DONG PROVINCE GENERAL HOSPITAL
EU-ME1 | 1543567 | 30 - 4 HOSPITAL
TJF-Q180V | 2923225 | KHANH HOA PROVINCE GENERAL HOSPITAL
GIF-H170 | 2929194 | 175 MILITARY HOSPITAL
GIF-XP170N | 2132010 | 175 MILITARY HOSPITAL
CF-H170I | 2711166 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CV-190 | 7050194 | CA MAU GENERAL HOSPITAL
CLV-190 | 7670938 | DONG NAI 2 HOSPITAL JOINT STOCK COMPANY
GIF-H170 | 2031855 | VICTORIA CLINIC
CF-H170I | 2922455 | MEDLATEC GENERAL HOSPITAL
GIF-EZ1500(E1) | 2102558 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
GIF-EZ1500(E1) | 2102541 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
GIF-H170 | 2828509 | VICTORIA CLINIC
CF-H170I | 2023315 | YEN BAI PROVINCIAL GENERAL HOSPITAL
CF-H170I | 2023311 | YEN BAI PROVINCIAL GENERAL HOSPITAL
TJF-Q180V | 2203406 | CITY INTERNATIONAL HOSPITAL
GIF-H180 | 2417946 | BECAMEX INTERNATIONAL HOSPITAL
WM-NP1 | 21132618 | COLUMBIA ASIA BINH DUONG INTERNATIONAL HOSPITAL
CF-HQ190I | 2962543 | DIGESTIVE & LIVER CENTER
CF-H170I | 2234749 | HANOI HOSPITAL OF TRADITIONAL MEDICAL
GIF-H170 | 2613507 | 108 MILITARY CENTRAL HOSPITAL
CF-H170I | 2821757 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-H190 | 2628461 | VINMEC NHA TRANG INTERNATIONAL GENERAL HOSPITAL
GIF-Q150 | 2002954 | 108 MILITARY CENTRAL HOSPITAL
GIF-H170 | 2725841 | HIEP HOA DISTRICT HOSPITAL
GIF-HQ190 | 2066523 | Phuong Chau International Hospital Soc Trang
ESG-400 | B000223 | TAN PHU HOSPITAL
CF-Q180AI | 2302676 | THONG NHAT HOSPITAL
CF-H170I | 2822154 | VICTORIA CLINIC
CF-H170I | 2023008 | HOAN MY THU DUC INTERNATIONAL GENERAL HOSPITAL
WA50373B | 801072 | UNIVERSITY MEDICAL CENTER HCMC
WA29042A | 800563 | UNIVERSITY MEDICAL CENTER HCMC
CF-Q150I | 2601850 | 108 MILITARY CENTRAL HOSPITAL
GIF-H170 | 2827449 | HOP LUC GENERAL HOSPITAL
GIF-H170 | 2827004 | MINISTRY OF PUBLIC SECURITY 199 HOSPITAL
UHI-4 | 7159664 | HUE CENTRAL HOSPITAL
CF-H170I | 2610819 | 108 MILITARY CENTRAL HOSPITAL
CV-180 | 7009462 | GIA DINH PEOPLE HOSPITAL
CLV-180 | 7018376 | GIA DINH PEOPLE HOSPITAL
CF-H170I | 2610820 | 108 MILITARY CENTRAL HOSPITAL
CHF-V(STERRAD NX) | 2962049 | HA DONG GENERAL HOSPITAL
BF-P190 | 2013823 | FV HOSPITAL
CLV-S190 | 7418103 | SAIGON ENT HOSPITAL
CV-170 | 7545694 | PHU YEN GENERAL HOSPITAL
GIF-Q180 | 2002313 | HA TRUNG GENERAL HOSPITAL
GIF-HQ190 | 2511187 | VINMEC NHA TRANG INTERNATIONAL GENERAL HOSPITAL
GIF-H170 | 2033368 | TAN PHU HOSPITAL
CF-HQ190I | 2062916 | CHO RAY HOSPITAL
GIF-HQ190 | 2065702 | CHO RAY HOSPITAL
CF-H170I | 2822211 | 175 MILITARY HOSPITAL
GIF-H190 | 2958985 | PHUONG DONG GENERAL HOSPITAL
GIF-H170 | 2828411 | VICTORIA CLINIC
GIF-Q150 | 2721360 | BECAMEX INTERNATIONAL HOSPITAL
WA2T430A | 819228 | HUE CENTRAL HOSPITAL
CF-H190I | 2842568 | PHUONG DONG GENERAL HOSPITAL
CF-Q150I | 2501825 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CF-HQ190I | 2862393 | YERSIN INTERNATIONAL CLINIC
TJF-Q180V | 2507503 | VINMEC CENTRAL PARK HOSPITAL
GIF-HQ190 | 2858928 | 175 MILITARY HOSPITAL
GIF-HQ190 | 2746055 | 175 MILITARY HOSPITAL
GIF-Q180 | 2705935 | 175 MILITARY HOSPITAL
GIF-H170 | 2726042 | 108 MILITARY CENTRAL HOSPITAL
GIF-HQ190 | 2064817 | 108 MILITARY CENTRAL HOSPITAL
CH-S400-XZ-EB | 7823377 | HUE CENTRAL HOSPITAL
CF-H170I | 2023005 | HOAN MY THU DUC INTERNATIONAL GENERAL HOSPITAL
SIF-Q180 | 2422533 | 108 MILITARY CENTRAL HOSPITAL
A70942A | 800312 | VICTORIA CLINIC
CF-H170I | 2922761 | 108 MILITARY CENTRAL HOSPITAL
CF-H170I | 2922921 | 108 MILITARY CENTRAL HOSPITAL
CF-H170I | 2922925 | 108 MILITARY CENTRAL HOSPITAL
MAJ-1597 | 1503950 | CHO RAY HOSPITAL
CLV-190 | 7777309 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
OER-AW | 2732400 | VINMEC CENTRAL PARK HOSPITAL
TJF-Q180V | 2134882 | THONG NHAT HOSPITAL
CF-Q150I | 2000448 | 108 MILITARY CENTRAL HOSPITAL
GIF-H190 | 2164818 | BAI CHAY HOSPITAL
GIF-Q180 | 2809320 | 108 MILITARY CENTRAL HOSPITAL
GIF-Q150 | 2317679 | 108 MILITARY CENTRAL HOSPITAL
GIF-Q150 | 2721456 | TAN PHU HOSPITAL
GIF-Q180 | 2203602 | TAN PHU HOSPITAL
GIF-LV1(PAL) | 1400897 | SAIGON - ITO HOSPITAL
CF-EZ1500DI(E1) | 2201470 | TAM ANH HOSPITAL- HCMC BRANCH
MAJ-1430 | 7935438 | NGUYEN TRAI HOSPITAL
CV-180 (PAL) | 7007373 | CAI LAY REGIONAL GENERAL HOSPITAL
CF-H170I | 2611094 | CHIEM HOA GENERAL HOSPITAL
CF-Q150I | 2000355 | HA TRUNG GENERAL HOSPITAL
TJF-Q180V | 2720793 | VINMEC DA NANG INTERNATIONAL GENERAL HOSPITAL
GIF-V70-PAL | 2411436 | HOC MON REGION GENERAL HOSPITAL
CV-170 | 7646347 | TAY NINH PROVINCE GENERAL HOSPITAL
GIF-H170 | 2827006 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CF-H170I | 2123727 | THAI NGUYEN CENTRE GENERAL HOSPITAL
BF-Q170 | 2700417 | VIET XO FRIENDSHIP HOSPITAL
GIF-H170 | 2401334 | SAIGON ENT HOSPITAL
GIF-Q150 | 2418932 | SAIGON ENT HOSPITAL
CLV-180 | 7112332 | UNIVERSITY MEDICAL CENTER HCMC 2
A70960A | 800315 | CAREPLUS CLINIC NO.2
TJF-Q180V | 2934220 | TRIEU AN HOSPITAL
BF-Q170 | 2001029 | PHAM NGOC THACH HOSPITAL
BF-P190 | 2812507 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
GIF-Q180 | 2900425 | HA TINH GENERAL HOSPITAL
CF-H170I | 2234596 | CU CHI REGIONAL GENERAL HOSPITAL
TJF-160VR | 2902366 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
TJF-160VR | 2500547 | 175 MILITARY HOSPITAL
ENF-P4 | 2156470 | EAR NOSE THROAT (ENT) HOSPITAL HCMC
ENF-P4 | 2156483 | EAR NOSE THROAT (ENT) HOSPITAL HCMC
BF-1T150(STERRAD NX) | 2751776 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
GIF-H170 | 2930351 | HAI PHONG INTERNATIONAL HOSPITAL
CF-H170I | 2922680 | HAI PHONG INTERNATIONAL HOSPITAL
GIF-HQ190 | 2066472 | HUNG HA GENERAL HOSPITAL
TJF-Q170V | 2200448 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-XP170N | 2400411 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-HQ190 | 2859413 | HOAN MY SAIGON HOSPITAL
CV-190 | 7749385 | BINH DINH GENERAL HOSPITAL JOINT STOCK COMPANY
TJF-Q170V | 2300587 | BACH MAI HOSPITAL
CF-Q150I | 2600010 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CF-H170I | 2610815 | 108 MILITARY CENTRAL HOSPITAL
TJF-150 | 2800087 | BINH AN HOSPITAL
GIF-1TH190 | 2942860 | FV HOSPITAL
CF-V70I-PAL | 2410378 | HOC MON REGION GENERAL HOSPITAL
GIF-H170 | 2134068 | EASTERN ARMY AND PEOPLE HOSPITAL
GIF-Q150 | 2104641 | QUYHOA NATIONAL LEPROSY DERMATOLOGY HOSPITAL
MAJ-1430 | 7112408 | QUYHOA NATIONAL LEPROSY DERMATOLOGY HOSPITAL
CLV-190 | 7776645 | BINH CHANH HOSPITAL
CF-HQ190I | 2862178 | VIET XO FRIENDSHIP HOSPITAL
WA2T430A | 816731 | NAM SAI GON INTERNATIONAL GENERAL HOSPITAL
GIF-H190 | 2368723 | KHANH HOA PROVINCE GENERAL HOSPITAL
GIF-Q150 | 2104251 | UNIVERSITY MEDICAL CENTER HCMC 2
GIF-H170 | 2134103 | UNIVERSITY MEDICAL CENTER HCMC 2
GIF-EZ1500(E1) | 2001568 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GF-UCT180 | 1621408 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
TJF-Q180V | 2034418 | HOAN MY SAIGON HOSPITAL
GIF-HQ190 | 2168769 | FV HOSPITAL
ENF-XP | 2926322 | NATIONAL HOSPITAL OF ODONTO - STOMATOLOGY - HO CHI MINH CITY (NHOS-HCMC)
BF-H190 | 2026967 | BECAMEX INTERNATIONAL HOSPITAL
OER-AW | 2033224 | VICTORIA CLINIC
CF-H170I | 2822238 | VIET NAM NATIONAL INSTITUTE OF MARITIME MEDICINE
GIF-HQ190 | 2168610 | CAN THO UNIVERSITY OF MEDICINE AND PHARMACY HOSPITAL
GIF-Q150 | 2821677 | THU DUC REGIONAL GENERAL HOSPITAL
CF-H180AI | 2403076 | BECAMEX INTERNATIONAL HOSPITAL
CLV-180 | 7215122 | CITY INTERNATIONAL HOSPITAL
URF-V(STERRAD NX) | 2661591 | BAC GIANG PROVINCE GENERAL HOSPITAL
A37025A | 705855 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
LTF-S300-10-3D | 2000991 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CYF-5A(STERRAD NX) | 2263349 | BINH DAN HOSPITAL
UHI-3 | 1707519 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CHF-V(STERRAD NX) | 2962055 | CU CHI REGIONAL GENERAL HOSPITAL
MAJ-1430 | 7218241 | CITY INTERNATIONAL HOSPITAL
GIF-EZ1500(E1) | 2102953 | TAM ANH HOSPITAL- HCMC BRANCH
GIF-EZ1500(E1) | 2304574 | TAM ANH HOSPITAL- HCMC BRANCH
GIF-HQ190 | 2612424 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
BF-1TQ170 | 2101423 | PHAM NGOC THACH HOSPITAL
OTV-SI | 7556849 | CAREPLUS CLINIC NO.2
GIF-Q150 | 2721450 | GIA DINH PEOPLE HOSPITAL
CF-EZ1500DI(E1) | 2201397 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
GIF-HQ190 | 2064994 | HUNG HA GENERAL HOSPITAL
GIF-EZ1500(E1) | 2203071 | 91 PHAM HUU CHI CLINIC
GIF-H170 | 2033149 | DOCTOR CHUNG CLINIC
MAJ-1430 | 7514457 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CLV-190 | 7204295 | HOAN MY BINH PHUOC HOSPITAL
CF-HQ190L | 2084022 | CA MAU GENERAL HOSPITAL
CV-170 | 7545121 | TRIEU AN HOSPITAL
GIF-EZ1500(E1) | 2304617 | TAM ANH HOSPITAL- HCMC BRANCH
GIF-H190 | 2164811 | BAI CHAY HOSPITAL
GIF-Q150 | 2002813 | HUNG YEN GENERAL HOSPITAL
TJF-Q180V | 2934204 | HOAN MY THU DUC INTERNATIONAL GENERAL HOSPITAL
ESG-400 | B011126 | HUE CENTRAL HOSPITAL - BRANCH 2
CLV-190 | 7775651 | SAI GON TORANOMON CLINIC
GIF-H170 | 2828558 | VIET - GERMANY AMITY HOSPITAL
GIF-H170 | 2033109 | QUE VO GENERAL HOSPITAL
GIF-H190 | 2856438 | HA NOI FRENCH HOSPITAL
GIF-H170 | 2930450 | 108 MILITARY CENTRAL HOSPITAL
CLV-180 | 7215146 | TAM BINH DISTRICT GENERAL HOSPITAL
MAJ-1430 | 7218706 | TAM BINH DISTRICT GENERAL HOSPITAL
CF-H170I | 2023427 | VIET XO FRIENDSHIP HOSPITAL
CV-V1 | 7830979 | LE VAN VIET HOSPITAL
MAJ-554 | 7707164 | SAIGON FV CLINIC
CF-EZ1500DL(E1) | 2201709 | TAM ANH HOSPITAL- HCMC BRANCH
GIF-EZ1500(E1) | 2203313 | TAM ANH HOSPITAL- HCMC BRANCH
GIF-H170 | 2929883 | THU DUC CITY HOSPITAL
CF-H170I | 2124079 | DONG ANH GENERAL HOSPITAL
CHF-V2 | 2000133 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
GIF-EZ1500(E1) | 2102948 | TAM ANH HOSPITAL- HCMC BRANCH
GIF-HQ190 | 2859384 | HOAN MY SAIGON HOSPITAL
GIF-H170 | 2400607 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
BF-H190 | 2026897 | 175 MILITARY HOSPITAL
GIF-Q150 | 2003002 | DISTRICT 8 HOSPITAL
BF-UC190F | 7211789 | UNIVERSITY MEDICAL CENTER HCMC
BF-UC190F | 7900307 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CF-H170I | 2610818 | GIA BINH DISTRICT GENERAL HOSPITAL
CF-H170I | 2922369 | GIA AN 115 HOSPITAL
BF-P150(STERRAD NX) | 2540963 | THANH NHAN HOSPITAL
GIF-H170 | 2501902 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-Q180 | 2203560 | QUANG NAM CENTRAL GENERAL HOSPITAL
GIF-H170 | 2614451 | LE VAN THINH HOSPITAL
CF-H170I | 2123758 | TAN PHU HOSPITAL
GIF-HQ190 | 2613016 | HOAN MY SAIGON HOSPITAL
GIF-HQ190 | 2746545 | TRUNG VUONG HOSPITAL
TJF-Q180V | 2507333 | TAY NGUYEN REGIONAL GENERAL HOSPITAL
GIF-H190 | 2265614 | PUBLIC SECURITY MINISTRY 19/8 HOSPITAL
CF-H170I | 2023201 | DISTRICT 1 HOSPITAL
GIF-H170 | 2826957 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-H170 | 2726298 | P & T GENERAL HOSPITAL NO.1
CF-H170I | 2610730 | QUE VO GENERAL HOSPITAL
CH-S190-08-LB | 7918021 | BINH DAN HOSPITAL
GIF-H170 | 2031502 | HOAN MY THU DUC INTERNATIONAL GENERAL HOSPITAL
BF-MP190F | 2700090 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-H170 | 2338716 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-XP190N | 2722608 | FV HOSPITAL
CF-H170I | 2123713 | NAM SAI GON INTERNATIONAL GENERAL HOSPITAL
GF-UCT180 | 1021547 | CHO RAY HOSPITAL
ESG-400 | B008937 | HUNG VUONG GIA LAI HOSPITAL
EU-ME2 | 7015980 | CHO RAY HOSPITAL
GIF-XP170N | 2931741 | E HOSPITAL
ENF-V3 | 2743230 | PEDIATRIC HOSPITAL NO. 2
GIF-H170 | 2613647 | SON TAY DISTRICT GENERAL HOSPITAL
CF-H170I | 2023084 | ONCOLOGY HOSPITAL ( BRANCH 2 HCM )
GIF-H170 | 2338376 | BINH THANH HOSPITAL
CF-H170I | 2821700 | DA NANG HOSPITAL
CF-HQ190I | 2862112 | DA NANG HOSPITAL
GIF-EZ1500(E1) | 2203886 | DA NANG HOSPITAL
CF-H170I | 2821864 | NGUYEN TRI PHUONG HOSPITAL
GIF-HQ190 | 2745630 | 91 PHAM HUU CHI CLINIC
CF-H170I | 2335059 | NINH BINH GENERAL HOSPITAL
CF-H170I | 2234611 | DOCTOR CHUNG CLINIC
CLH-SC | 7132144 | TU DU HOSPITAL
CF-HQ190L | 2394971 | BINH DAN HOSPITAL
CF-HQ190L | 2394950 | BINH DAN HOSPITAL
GIF-H170 | 2032261 | BINH DAN HOSPITAL
CV-V1 | 7031881 | CHU SE DISTRICT GENERAL HOSPITAL
GIF-LV1(PAL) | 2012547 | CHU SE DISTRICT GENERAL HOSPITAL
CF-HQ190I | 2062782 | 108 MILITARY CENTRAL HOSPITAL
GIF-XP170N | 2500974 | 108 MILITARY CENTRAL HOSPITAL
GIF-XP190N | 2722618 | FV HOSPITAL
GIF-EZ1500(E1) | 2306105 | IMED MEDICAL EQUIPMENT Co., Ltd
CF-EZ1500DI(E1) | 2301919 | IMED MEDICAL EQUIPMENT Co., Ltd
CF-H170I | 2234567 | HAI PHONG INTERNATIONAL HOSPITAL
GIF-H170 | 2930344 | HAI PHONG INTERNATIONAL HOSPITAL
ENF-V3 | 2743237 | PEDIATRIC HOSPITAL NO. 2
GIF-H170 | 2236731 | HUNG VUONG - PHU THO GENERAL HOSPITAL
CLV-S400 | 7711411 | 175 MILITARY HOSPITAL
ENF-VT2 | 2240678 | ONCOLOGY HOSPITAL ( BRANCH 2 HCM )
CF-H170I | 2023090 | ONCOLOGY HOSPITAL ( BRANCH 2 HCM )
GIF-H190 | 2629145 | FV HOSPITAL
CF-H190I | 2501111 | 91 PHAM HUU CHI CLINIC
GIF-H190 | 2369531 | NGUYEN TRAI HOSPITAL
CF-HQ190I | 2163150 | PRIMA PXL MEDICAL CENTER
GIF-EZ1500(E1) | 2102951 | TAM ANH HOSPITAL- HCMC BRANCH
GIF-H190 | 2629181 | VINMEC HA LONG INTERNATIONAL HOSPITAL
GIF-HQ190 | 2064782 | 108 MILITARY CENTRAL HOSPITAL
GIF-HQ190 | 2373839 | HUNG HA PHO NOI CLINIC
BF-1T150(STERRAD NX) | 2751801 | 30 - 4 HOSPITAL
GIF-H170 | 2338480 | TAN BINH HOSPITAL
GIF-EZ1500(E1) | 2203307 | TAM ANH HOSPITAL- HCMC BRANCH
CF-HQ190I | 2511038 | E HOSPITAL
GIF-H170 | 2827552 | AMERICAN INTERNATIONAL HOSPITAL
GIF-H170 | 2338391 | HANOI INTERNATIONAL HOSPITAL - BAC GIANG
GIF-H170 | 2237139 | DONG DO HOSPITAL
CF-H170I | 2334963 | HUNG VUONG GIA LAI HOSPITAL
CF-H170I | 2334927 | HUNG VUONG GIA LAI HOSPITAL
CF-H170I | 2023360 | HAI PHONG INTERNATIONAL HOSPITAL
CF-H170I | 2822209 | HAI PHONG INTERNATIONAL HOSPITAL
CF-Q150I | 2301395 | THANH NHAN HOSPITAL
OTV-S7PRO | 7579316 | BINH DINH PROVINCE GENERAL HOSPITAL
GIF-H190 | 2414468 | FV HOSPITAL
CF-EZ1500DI(E1) | 2201465 | TAM ANH HOSPITAL- HCMC BRANCH
GIF-H180 | 2900802 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
UHI-4 | 7161000 | NAM SAI GON INTERNATIONAL GENERAL HOSPITAL
GIF-H190 | 2164597 | HA GIANG PROVINCE GENERAL HOSPITAL
GIF-HQ190 | 2066140 | THU DUC REGIONAL GENERAL HOSPITAL
CF-H190I | 2143950 | THONG NHAT DONG NAI GENERAL HOSPITAL
GIF-H190 | 2162981 | THONG NHAT DONG NAI GENERAL HOSPITAL
CV-170 | 7985968 | THU DUC CITY HOSPITAL
GIF-HQ190 | 2859007 | BAI CHAY HOSPITAL
MAJ-1558 | 7902995 | GIA DINH PEOPLE HOSPITAL
CF-Q150I | 2900250 | NATIONAL INSTITUTE OF HEMATOLOGY AND BLOOD TRANSFUSION
CF-Q150I | 2711931 | NATIONAL INSTITUTE OF HEMATOLOGY AND BLOOD TRANSFUSION
BF-P60(STERRAD NX) | 2454980 | LUNG & TUBERCULOSIS HOSPITAL OF LANG SON PROVINCE
BF-P60(STERRAD NX) | 2043017 | LUNG & TUBERCULOSIS HOSPITAL OF LANG SON PROVINCE
CF-Q150I | 2900196 | TAN UYEN GENERAL HOSPITAL
GIF-Q150 | 2901304 | TAN UYEN GENERAL HOSPITAL
GIF-H190 | 2960153 | YERSIN INTERNATIONAL CLINIC
BF-Q170 | 2201769 | HA NOI MEDICAL UNIVERSITY HOSPITAL
GIF-EZ1500(E1) | 2305051 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-HQ190 | 2612451 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-H170 | 2033331 | TAM TRI DA NANG GENERAL HOSPITAL
GIF-HQ190 | 2066157 | THU DUC REGIONAL GENERAL HOSPITAL
CF-H170I | 2123989 | AGRICULTURE HOSPITAL
GIF-Q150 | 2519926 | E HOSPITAL
GIF-H170 | 2614643 | 125 THAI THINH CLINIC
CV-170 | 7646069 | BINH DINH PROVINCE GENERAL HOSPITAL
GIF-H170 | 2338389 | BACH MAI HOSPITAL
GIF-H170 | 2236695 | HAI PHONG INTERNATIONAL HOSPITAL
GIF-HQ190 | 2613007 | HOAN MY SAIGON HOSPITAL
GIF-HQ190 | 2746136 | DAI PHUOC GENERAL CLINIC
GIF-EZ1500(E1) | 2102547 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-H190 | 2527257 | 91 PHAM HUU CHI CLINIC
CF-EZ1500DI(E1) | 2101323 | 91 PHAM HUU CHI CLINIC
GIF-HQ190 | 2634713 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CLV-180 | 7215181 | CAM LAM GENERAL HOSPITAL
GIF-Q150 | 2317704 | CAM LAM GENERAL HOSPITAL
GIF-H170 | 2929419 | MEDLATEC GENERAL HOSPITAL
GIF-H190 | 2855304 | PEDIATRIC HOSPITAL NO. 1
CF-HQ190I | 2263244 | HONG NGOC GROUP - PHUC TRUONG MINH BRANCH
TJF-Q180V | 2506152 | UNIVERSITY MEDICAL CENTER HCMC
CF-H170I | 2023599 | HAI PHONG INTERNATIONAL HOSPITAL
GIF-Q150 | 2002833 | THAI NGUYEN CENTRE GENERAL HOSPITAL
GIF-H180 | 2900809 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CLV-190 | 7882311 | KHANH HOA PROVINCE GENERAL HOSPITAL
ESG-400 | 16069W14001 | BINH DINH PROVINCE GENERAL HOSPITAL
GIF-H170 | 2614346 | VICTORIA CLINIC
CF-HQ190I | 2862068 | FV HOSPITAL
GIF-HQ190 | 2169410 | UNIVERSITY MEDICAL CENTER HCMC
GIF-HQ190 | 2169327 | UNIVERSITY MEDICAL CENTER HCMC
GIF-HQ190 | 2963839 | UNIVERSITY MEDICAL CENTER HCMC
GIF-HQ190 | 2858923 | UNIVERSITY MEDICAL CENTER HCMC
CF-HQ190I | 2163165 | UNIVERSITY MEDICAL CENTER HCMC
GIF-H170 | 2613802 | BAC GIANG ONCOLOGY HOSPITAL
GIF-H170 | 2614084 | TAM DUONG DISTRICT HOSPITAL
CV-170 | 7544159 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-H190 | 2369125 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
BF-1TQ170 | 2600269 | FV HOSPITAL
CF-EZ1500DI(E1) | 2101281 | DIGESTIVE & LIVER CENTER
GIF-Q150 | 2901497 | NATIONAL INSTITUTE OF HEMATOLOGY AND BLOOD TRANSFUSION
GIF-H190 | 2470339 | PEDIATRIC HOSPITAL NO. 1
GIF-H170 | 2715193 | GIA DINH PEOPLE HOSPITAL
BF-Q170 | 2400083 | BAC GIANG LUNG AND TUBERCULOSIS HOSPITAL
GIF-H170 | 2930431 | 354 MILITARY HOSPITAL
GIF-H170 | 2930441 | 354 MILITARY HOSPITAL
CF-H170I | 2822142 | 354 MILITARY HOSPITAL
GIF-H170 | 2827454 | 354 MILITARY HOSPITAL
GIF-H170 | 2930436 | 354 MILITARY HOSPITAL
CF-H170I | 2822215 | 354 MILITARY HOSPITAL
GIF-H170 | 2828563 | 354 MILITARY HOSPITAL
OTV-S200 | 7355486 | NAM SAI GON INTERNATIONAL GENERAL HOSPITAL
CF-H190I | 2043868 | E HOSPITAL
GIF-EZ1500(E1) | 2102524 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
UHI-4 | 7160887 | TAM ANH HOSPITAL- HCMC BRANCH
GIF-H170 | 2614285 | CU CHI DISTRICT HOSPITAL
GIF-Q180 | 2625067 | PEDIATRIC HOSPITAL NO. 2
GIF-H170 | 2501730 | HIEP HOA DISTRICT HOSPITAL
MAJ-1430 | 7833292 | XUYEN A GENERAL HOSPITAL - VINH LONG
BF-XP190 | 2600702 | PEDIATRIC HOSPITAL NO. 2
CLV-190 | 7882308 | KHANH HOA PROVINCE GENERAL HOSPITAL
GIF-H170 | 2135021 | HAI PHONG INTERNATIONAL HOSPITAL - VINH BAO
GIF-EZ1500(E1) | 2203894 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CF-EZ1500DI(E1) | 2000454 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CV-180 | 7009041 | DAK TO DISTRICT MEDICAL CENTER
CLV-180 | 7017794 | DAK TO DISTRICT MEDICAL CENTER
CF-EZ1500DI(E1) | 2000612 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CHF-V(STERRAD NX) | 2861836 | THONG NHAT DONG NAI GENERAL HOSPITAL
GIF-H170 | 2031501 | HOAN MY THU DUC INTERNATIONAL GENERAL HOSPITAL
GIF-H170 | 2033006 | THAI NGUYEN INTERNATIONAL HOSPITAL
CF-HQ190I | 2962537 | PHUONG CHAU SA DEC INTERNATIONAL HOSPITAL
CF-EZ1500DI(E1) | 2101286 | TAM ANH HOSPITAL- HCMC BRANCH
TJF-Q190V | 2431550 | CHO RAY HOSPITAL
GIF-EZ1500(E1) | 2304619 | TAM ANH HOSPITAL- HCMC BRANCH
GIF-HQ190 | 2065701 | CHO RAY HOSPITAL
CF-HQ190I | 2062948 | CHO RAY HOSPITAL
ESG-400 | B006289 | HUNG VUONG HOSPITAL
BF-1TQ170 | 2101630 | 175 MILITARY HOSPITAL
CF-H170I | 2123825 | BAC THANG LONG LUC NGAN CLINIC
TD-20 | 102722 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-EZ1500(E1) | 2203893 | 22-12 HOSPITAL
UHI-4 | 7741003 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
UHI-4 | 7210461 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CF-H170I | 2922536 | THU DUC CITY HOSPITAL
BF-Q170 | 2700342 | YEN BAI PROVINCIAL GENERAL HOSPITAL
GIF-H170 | 2338717 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-Q150 | 2519644 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-Q150 | 2002311 | CAN THO UNIVERSITY OF MEDICINE AND PHARMACY HOSPITAL
UHI-4 | 7638140 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
BF-H190 | 2026706 | MILITARY HOSPITAL NO.103
CF-H170I | 2023365 | HAI PHONG INTERNATIONAL HOSPITAL
GIF-H170 | 2929573 | MEDLATEC CLINIC (THANH XUAN branch)
UHI-4 | 7740991 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
TJF-160VR | 2400031 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CV-170 | 7987318 | 175 MILITARY HOSPITAL
GIF-Q150 | 2922394 | NINH BINH GENERAL HOSPITAL
TJF-Q180V | 2721710 | TRUNG VUONG HOSPITAL
CF-HQ190I | 2631360 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
GIF-HQ190 | 2375043 | UNIVERSITY MEDICAL CENTER HCMC
CF-H190I | 2244859 | KHANH HOA PROVINCE GENERAL HOSPITAL
UHI-4 | 7323036 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
ESG-400 | B002291 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
ESG-400 | B002624 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CF-H170I | 2922994 | ONCOLOGY HOSPITAL
CF-H170I | 2922982 | ONCOLOGY HOSPITAL
UHI-4 | 7946224 | HUE CENTRAL HOSPITAL
TJF-Q190V | 2124352 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
BF-Q170 | 2201814 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CF-HQ190I | 2761913 | 91 PHAM HUU CHI CLINIC
CF-EZ1500DI(E1) | 2201398 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
UHI-4 | 7537358 | LAM DONG PROVINCE GENERAL HOSPITAL
GIF-H170 | 2236583 | THAI HOA GENERAL CLINIC
GIF-EZ1500(E1) | 2304624 | TAM ANH HOSPITAL- HCMC BRANCH
GIF-EZ1500(E1) | 2304656 | TAM ANH HOSPITAL- HCMC BRANCH
GIF-EZ1500(E1) | 2102998 | TAM TRI DA NANG GENERAL HOSPITAL
GIF-EZ1500(E1) | 2305108 | Y VIET JOINT STOCK COMPANY - UNIVERSITY MEDICAL HOSPITAL CLINIC 1
GIF-HQ190 | 2961790 | PHUONG CHAU SA DEC INTERNATIONAL HOSPITAL
GIF-HQ190 | 2373012 | UNIVERSITY MEDICAL CENTER HCMC
GIF-HQ190 | 2613380 | VINMEC HA LONG INTERNATIONAL HOSPITAL
CF-HQ190I | 2761910 | VINMEC HA LONG INTERNATIONAL HOSPITAL
CF-H170I | 2510654 | THU DUC CITY HOSPITAL
CF-H170I | 2721627 | BINH CHANH HOSPITAL
MAJ-1430 | 7937155 | 175 MILITARY HOSPITAL
CLV-190 | 7087705 | THONG NHAT HOSPITAL
BF-Q170 | 2001026 | VIET NAM-THUY DIEN UONG BI HOSPITAL
CF-HQ190I | 2062975 | HA NOI FRENCH HOSPITAL
MAJ-1430 | 7514277 | XUYEN A GENERAL HOSPITAL - VINH LONG
CF-EZ1500DI(E1) | 2201395 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
BF-1TQ170 | 2402233 | PHAM NGOC THACH HOSPITAL
WA4KL130 | 803070 | HUE CENTRAL HOSPITAL
CLH-SC | 7626687 | TU DU HOSPITAL
BF-1TQ170 | 2101434 | 175 MILITARY HOSPITAL
BF-1TQ170 | 2402247 | PHAM NGOC THACH HOSPITAL
CF-H170I | 2023195 | FAMILY HEALTH CARE JSC
CF-EZ1500DI(E1) | 2000650 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
UHI-4 | 7649015 | TAN PHU HOSPITAL
UCR | 7835062 | VICTORIA CLINIC
MAJ-1430 | 7015141 | 175 MILITARY HOSPITAL
CV-170 | 7316730 | YUN PA DISTRICT GENERAL HOSPITAL
GIF-H170 | 2929739 | HOP LUC GENERAL HOSPITAL
TC-E400 | 7303034 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-H170 | 2930321 | THAIBINH HOSPITAL.,CO LTD
GIF-H170 | 2440723 | BINH DINH PROVINCE GENERAL HOSPITAL
SPL-G | CG0175 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CH-S400-XZ-EB | 7331713 | PHUC HUNG GENERAL HOSPITAL
OTV-S400 | 7352000 | PHUC HUNG GENERAL HOSPITAL
WM-NP2 | 21303320 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
WM-NP2 | 21524885 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
ESG-400 | 14079W130005 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CV-1500 | 7134946 | DIGESTIVE & LIVER CENTER
CV-190 | 7841395 | PEDIATRIC HOSPITAL NO. 1
CF-H170I | 2234615 | DONG DO HOSPITAL
CF-H170I | 2922727 | 108 MILITARY CENTRAL HOSPITAL
GIF-HQ190 | 2859114 | BAI CHAY HOSPITAL
GIF-HQ190 | 2859196 | DAI PHUOC GENERAL CLINIC
TJF-Q170V | 2400662 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-Q150 | 2002395 | TAN PHUOC MEDICAL CENTER
CF-HQ190I | 2962575 | DIGESTIVE & LIVER CENTER
GIF-H190 | 2368998 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
BF-1TQ170 | 2302083 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
PCF-H190I | 2730538 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-H170 | 2338535 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CYF-VHA | 2400633 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-H170 | 2930437 | 108 MILITARY CENTRAL HOSPITAL
CH-S190-XZ-E | 7607021 | THU DUC CITY HOSPITAL
CV-190 | 7622040 | THU DUC CITY HOSPITAL
CLV-190 | 7669093 | THU DUC CITY HOSPITAL
UHI-4 | 7741516 | DISTRICT 6 HOSPITAL
CF-EZ1500DI(E1) | 2201457 | TAM ANH HOSPITAL- HCMC BRANCH
GIF-XP170N | 2232205 | LU VAN TRANG GENERAL CLINIC
CV-190 | 7050048 | AN GIANG CENTRAL GENERAL HOSPITAL
CHF-V2 | 2900097 | THONG NHAT HOSPITAL
OTV-S7PROH-HD-10E | 7543660 | HOAN MY SAIGON HOSPITAL
OTV-S7PROH-HD-12E | 7444608 | SAIGON ENT HOSPITAL
CF-H170I | 2023156 | HAI PHONG INTERNATIONAL HOSPITAL - VINH BAO
CF-H170I | 2124039 | HOP LUC GENERAL HOSPITAL
GIF-EZ1500(E1) | 2102545 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CF-HQ190I | 2862059 | VINMEC HAI PHONG INTERNATIONAL GENERAL HOSPITAL
CV-190 | 7629869 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-Q150 | 2721444 | THONG NHAT HOSPITAL
CF-H170I | 2822310 | QUANG NGAI GENERAL HOSPITAL
GIF-H170 | 2031760 | E HOSPITAL
OFP-2 | 22312762 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CF-HQ190I | 2631426 | VINMEC HA LONG INTERNATIONAL HOSPITAL
GIF-H170 | 2715644 | VICTORIA CLINIC
GIF-H180J | 2418218 | E HOSPITAL
GIF-H170 | 2134038 | MEDLATEC GENERAL HOSPITAL
GIF-Q180 | 2204001 | DA NANG C HOSPITAL
CF-HQ190I | 2062748 | DA NANG C HOSPITAL
UHI-4 | 7740981 | 175 MILITARY HOSPITAL
GIF-H170 | 2931453 | 108 MILITARY CENTRAL HOSPITAL
CF-H170I | 2500385 | HIEP HOA DISTRICT HOSPITAL
GIF-H190 | 2060693 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CF-H170I | 2922991 | ONCOLOGY HOSPITAL
GIF-H170 | 2614317 | SHINGMARK MEDICAL UNIVERSITY HOSPITAL
CF-HQ190L | 2394931 | BINH DAN HOSPITAL
WM-NP2 | 21995198 | CHO RAY HOSPITAL
GIF-H170 | 2235824 | PRIMA PXL MEDICAL CENTER
CF-EZ1500DI(E1) | 2201396 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CV-170 | 7882745 | UNIVERSITY MEDICAL CENTER HCMC
CF-H170I | 2335035 | HAI PHONG INTERNATIONAL HOSPITAL
ENF-V2(STERRAD NX) | 2861183 | FV HOSPITAL
OTV-S700 | 7301437 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
UHI-4 | 7945554 | CAU NGANG DISTRICT GENERAL HOSPITAL
A22001A | 609683 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
A22002A | 725269 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CF-EZ1500DI(E1) | 2201386 | DOCTOR CHECK CLINIC
CF-H170I | 2822067 | E HOSPITAL
GIF-Q150 | 2317684 | 108 MILITARY CENTRAL HOSPITAL
GIF-H170 | 2134582 | AGRICULTURE HOSPITAL
GIF-H170 | 2134114 | BAC THANG LONG LUC NGAN CLINIC
GIF-XP190N | 2336595 | PEDIATRIC HOSPITAL NO. 1
UHI-4 | 7058753 | TIEN GIANG CENTRAL GENERAL HOSPITAL
GIF-HQ190 | 2757276 | DIGESTIVE & LIVER CENTER
OFP-2 | 21994472 | VICTORIA CLINIC
GIF-H170 | 2133740 | TTH GENERAL HOSPITAL HA TINH
GIF-H170 | 2033413 | TTH GENERAL HOSPITAL HA TINH
GIF-H170 | 2033195 | THAI NGUYEN INTERNATIONAL HOSPITAL
A22001A | 641485 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
A22002A | 768899 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-XP190N | 2612238 | PEDIATRIC HOSPITAL NO. 2
GIF-EZ1500(E1) | 2102899 | TAM ANH HOSPITAL- HCMC BRANCH
CV-190 | 7844272 | YERSIN INTERNATIONAL CLINIC
GIF-HQ190 | 2859121 | DOLIFE INTERNATIONAL HOSPITAL
GIF-H170 | 2033288 | 175 MILITARY HOSPITAL
GIF-Q180 | 2102394 | PEDIATRIC HOSPITAL NO. 2
CF-HQ190I | 2363468 | LAC VIET FRIENDSHIP HOSPITAL
CF-H170I | 2023424 | LAM DONG II GENERAL HOSPITAL
GIF-EZ1500(E1) | 2306105 | THIEN HANH GENERAL HOSPITAL
CLV-180 | 7808576 | GIA DINH PEOPLE HOSPITAL
CF-Q150I | 2800091 | HONG DUC GENERAL HOSPITAL
GIF-Q150 | 2821527 | HONG DUC GENERAL HOSPITAL
UHI-4 | 7843163 | TIEN GIANG CENTRAL GENERAL HOSPITAL
ENF-VT2(STERRAD NX) | 2852005 | 175 MILITARY HOSPITAL
CF-H170I | 2023553 | VUNG TAU HOSPITAL
ENF-VH2 | 2200667 | VINMEC CENTRAL PARK HOSPITAL
UHI-4 | 7843704 | XUYEN A GENERAL HOSPITAL - VINH LONG
GIF-HQ190 | 2859372 | HOAN MY SAIGON HOSPITAL
A22001A | 598554 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CF-HQ190I | 2463619 | UNIVERSITY MEDICAL CENTER HCMC
CF-HQ190I | 2463618 | UNIVERSITY MEDICAL CENTER HCMC
GIF-H170 | 2726296 | ME LINH GENERAL HOSPITAL
GIF-H170 | 2031619 | BAC KAN PROVINCIAL GENERAL HOSPITAL
GIF-H180 | 2417899 | BECAMEX INTERNATIONAL HOSPITAL
CV-190 | 7211780 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
GIF-H170 | 2931039 | P & T GENERAL HOSPITAL NO.1
CF-H170I | 2721578 | P & T GENERAL HOSPITAL NO.1
GIF-H170 | 2613803 | P & T GENERAL HOSPITAL NO.1
GIF-H170 | 2931029 | P & T GENERAL HOSPITAL NO.1
CF-H170I | 2610776 | 30 - 4 HOSPITAL
GIF-H190 | 2369506 | NGUYEN TRAI HOSPITAL
UHI-4 | 7648802 | 17 MILITARY HOSPITAL
ESG-400 | B008527 | UNIVERSITY MEDICAL CENTER HCMC
USG-400 | 9015418 | UNIVERSITY MEDICAL CENTER HCMC
WA2T430A | 824958 | VIETNAM MEDICAL EQUIPMENT MATERIALS TRADING COMPANY LIMITED
GIF-H170 | 2338548 | HUNG VUONG GIA LAI HOSPITAL
CF-H170I | 2023557 | 175 MILITARY HOSPITAL
GIF-H170 | 2931023 | 175 MILITARY HOSPITAL
UHI-4 | 7843286 | DI LINH DISTRICT MEDICAL CENTRE
CF-H170I | 2023613 | TAM TRI DA NANG GENERAL HOSPITAL
CF-H170I | 2023298 | VIET TIEP HOSPITAL
ESG-400 | B002002 | DIEN KHANH DISTRICT MEDICAL CENTER
GIF-H190 | 2630434 | TAY NGUYEN REGIONAL GENERAL HOSPITAL
ENF-P4 | 7450636 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
WA2T430A | 824958 | BINH DAN HOSPITAL
CHF-V | 2651033 | PHU THO GENERAL HOSPITAL
CF-H170I | 2500386 | PHU THO GENERAL HOSPITAL
GIF-H170 | 2726219 | PHU THO GENERAL HOSPITAL
CF-LV1I(PAL) | 2010667 | COAL HOSPITAL
GIF-HQ190 | 2860860 | DIGESTIVE & LIVER CENTER
GIF-H170 | 2033256 | 175 MILITARY HOSPITAL
GIF-HQ190 | 2757266 | PUBLIC SECURITY MINISTRY 19/8 HOSPITAL
MAJ-1430 | 7214611 | TAN PHU HOSPITAL
WA2T430A | 824942 | BINH DAN HOSPITAL
CF-H170I | 2234776 | MINH ANH HOSPITAL
GIF-H170 | 2930382 | THAI NGUYEN INTERNATIONAL HOSPITAL
CF-H170I | 2234513 | THAI HOA GENERAL CLINIC
GIF-EZ1500(E1) | 2305002 | TRAN-LE TRADING AND PRODUCTION CO., LTD
GIF-EZ1500(E1) | 2305109 | TAM ANH GENERAL CLINIC - BRANCH OF TAM ANH HOSPITAL
CF-HQ190I | 2163110 | CAN THO UNIVERSITY OF MEDICINE AND PHARMACY HOSPITAL
UHI-4 | 7843294 | LAM HA DISTRICT MEDICAL CENTRE
GIF-Q180 | 2705964 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
BF-Q170 | 2302087 | PHAM NGOC THACH HOSPITAL
CF-H170I | 2611140 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
UHI-4 | 7537241 | THONG NHAT HOSPITAL
GIF-H170 | 2033190 | THAI NGUYEN INTERNATIONAL HOSPITAL
GIF-H170 | 2501639 | TIEN GIANG PROVINCE GENERAL HOSPITAL - SPECIAL TREATMENT AREA
GIF-H190 | 2471510 | MONG CAI HOSPITAL
TJF-Q190V | 2331527 | BINH DAN HOSPITAL
ESG-100 | 14085W500011 | SEAGEARO LONG AN HOSPITAL
GIF-H170 | 2931022 | 175 MILITARY HOSPITAL
BF-Q170 | 2001022 | VINH LONG GENERAL HOSPITAL
CF-HQ190L | 2637653 | BINH DAN HOSPITAL
GIF-H170 | 2440950 | TAM TRI NHA TRANG GENERAL HOSPITAL
BF-H190 | 2026904 | 175 MILITARY HOSPITAL
CF-EZ1500DI(E1) | 2201481 | PHUONG CHAU INTERNATIONAL HOSPITAL
GIF-EZ1500(E1) | 2305096 | NAM SAI GON INTERNATIONAL GENERAL HOSPITAL
CF-H170I | 2922564 | BINH THANH HOSPITAL
BF-UC190F | 7110874 | FV HOSPITAL
UHI-4 | 7843715 | HOAN MY SAIGON HOSPITAL
OTV-S7PRO | 7899571 | SON TRA DISTRICT MEDICAL CENTER
GIF-EZ1500(E1) | 2304876 | PHUONG CHAU INTERNATIONAL HOSPITAL
TJF-Q170V | 2300596 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
UHI-4 | 7740057 | QUANG NGAI OBSTETRICS AND PEDIATRICS HOSPITAL
GIF-HQ190 | 2475393 | CHO RAY HOSPITAL
WM-NP2 | 21605550 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CF-H170I | 2023181 | HAI PHONG INTERNATIONAL HOSPITAL - VINH BAO
OEV321UH | 7378103 | AN SINH HOSPITAL
CF-HQ190I | 2263341 | HOAN MY SAIGON HOSPITAL
GIF-EZ1500(E1) | 2203376 | TAM ANH HOSPITAL- HCMC BRANCH
CF-H170I | 2023602 | TTH GENERAL HOSPITAL HA TINH
GIF-H170 | 2613142 | YEN PHONG DISTRICT HOSPITAL
OER-AW | 2531544 | DONG NAI 2 HOSPITAL JOINT STOCK COMPANY
UHI-4 | 7263100 | VINH LONG GENERAL HOSPITAL
GIF-EZ1500(E1) | 2203887 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
ENF-V4 | 2200642 | HONG NGOC GROUP - PHUC TRUONG MINH BRANCH
CF-EZ1500DI(E1) | 2101283 | TAM ANH HOSPITAL- HCMC BRANCH
OTV-S7H-1N | 7403141 | PEDIATRIC HOSPITAL NO. 1
ENF-VH2 | 2900555 | 175 MILITARY HOSPITAL
CV-1500 | 7249063 | DAI PHUOC GENERAL CLINIC
GIF-H170 | 2340209 | K120 HOSPITAL
GIF-HQ190 | 2271590 | DAI PHUOC GENERAL CLINIC
GIF-H170 | 2033299 | 175 MILITARY HOSPITAL
CF-H170I | 2023605 | 175 MILITARY HOSPITAL
UHI-4 | 7945128 | KHANH HOA PROVINCE GENERAL HOSPITAL
GIF-H170 | 2340328 | 7A MILITARY HOSPITAL
CF-H170I | 2335316 | 7A MILITARY HOSPITAL
OFP-2 | 22329510 | DYM Medical Center Vietnam
GIF-H170 | 2440715 | BINH DINH PROVINCE GENERAL HOSPITAL
UHI-4 | 7649260 | PHU YEN GENERAL HOSPITAL
CF-H190I | 2345099 | VIETNAM MEDICAL EQUIPMENT MATERIALS TRADING COMPANY LIMITED
UHI-4 | 7843162 | TIEN GIANG OBSTETRICS & GYNAECOLOGY HOSPITAL
ESG-400 | B004116 | TIEN GIANG OBSTETRICS & GYNAECOLOGY HOSPITAL
GIF-HQ190 | 2475363 | CHO RAY HOSPITAL
CF-HQ190I | 2363473 | CHO RAY HOSPITAL
GIF-H170 | 2401239 | PHU THO GENERAL HOSPITAL
ENF-V3 | 2945146 | TTH GENERAL HOSPITAL HA TINH
UHI-4 | 7843283 | KBANG DISTRICT MEDICAL CENTER
GIF-H170 | 2032184 | LAP THACH HEALTH CENTRE
A70960A | 801174 | FAMILY MEDICAL PRACTICE HCMC CLINIC
BF-1T150 | 2700044 | HUE UNIVERSITY OF MEDICINE GENERAL HOSPITAL
GIF-XP170N | 2132016 | DYM Medical Center Vietnam
GIF-EZ1500(E1) | 2306773 | THIEN HANH GENERAL HOSPITAL
CF-EZ1500DI(E1) | 2201488 | Y VIET JOINT STOCK COMPANY - UNIVERSITY MEDICAL HOSPITAL CLINIC 1
UHI-4 | 7945970 | HUNG VUONG HOSPITAL
CF-H170I | 2234836 | PHUC YEN GENERAL HOSPITAL
CF-H170I | 2611120 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
UHI-4 | 7945122 | KHANH HOA PROVINCE GENERAL HOSPITAL
CF-H170I | 2500399 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
TJF-Q180V | 2934304 | NINH BINH GENERAL HOSPITAL
GIF-Q150 | 2519296 | QUANG NGAI GENERAL HOSPITAL
GIF-H170 | 2512918 | THU DUC CITY HOSPITAL
GIF-H170 | 2512924 | THU DUC CITY HOSPITAL
CF-H170I | 2922683 | THAIBINH HOSPITAL.,CO LTD
GIF-H170 | 2340283 | 7A MILITARY HOSPITAL
CF-H170I | 2335361 | 7A MILITARY HOSPITAL
WA4KL130 | 805958 | HUE CENTRAL HOSPITAL
GIF-HQ190 | 2475396 | CHO RAY HOSPITAL
GIF-HQ190 | 2475302 | CHO RAY HOSPITAL
MAJ-1558 | 7106454 | QUANG NAM CENTRAL GENERAL HOSPITAL
CF-H170I | 2500357 | TIEN GIANG PROVINCE GENERAL HOSPITAL - SPECIAL TREATMENT AREA
ESG-400 | B012022 | NGUYEN DINH CHIEU HOSPITAL
CF-H170I | 2822170 | HOP LUC INTERNATIONAL HOSPITAL - NGHI SON
GIF-Q150 | 2204671 | HA NAM PROVINCIAL GENERAL HOSPITAL
GIF-EZ1500(E1) | 2304615 | TAM ANH HOSPITAL- HCMC BRANCH
ESG-400 | B002216 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CF-HQ190I | 2651621 | ANH KHOA MEDICAL EQUIPMENT Co., LTD
GIF-H170 | 2929100 | QUANG NGAI GENERAL HOSPITAL
UHI-4 | 7945983 | HUNG VUONG HOSPITAL
ENF-VT2 | 2651424 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
TJF-Q180V | 2303542 | BAI CHAY HOSPITAL
CF-EZ1500DI(E1) | 2101287 | TAM ANH HOSPITAL- HCMC BRANCH
CF-H170L | 2440590 | GIA DINH PEOPLE HOSPITAL
ENF-P4 | 7450636 | EAR NOSE THROAT (ENT) HOSPITAL HCMC
CV-190 | 7749383 | 175 MILITARY HOSPITAL
UHI-4 | 7945118 | KHANH HOA PROVINCE GENERAL HOSPITAL
GIF-HQ190 | 2167533 | VINH LONG GENERAL HOSPITAL
CF-EZ1500DI(E1) | 2402230 | DONG DO HOSPITAL
CF-EZ1500DI(E1) | 2201452 | TAM ANH HOSPITAL- HCMC BRANCH
GIF-H170 | 2826753 | DA NANG HOSPITAL
CF-H170I | 2922685 | DA NANG HOSPITAL
GIF-EZ1500(E1) | 2203793 | DOCTOR CHECK CLINIC
GIF-H170 | 2930342 | THAIBINH HOSPITAL.,CO LTD
OTV-S200 | 7032159 | BINH DAN HOSPITAL
CF-EZ1500DI(E1) | 2201368 | YERSIN INTERNATIONAL CLINIC
CF-H170I | 2500366 | 108 MILITARY CENTRAL HOSPITAL
GIF-H170 | 2929267 | 108 MILITARY CENTRAL HOSPITAL
GIF-H170 | 2512763 | 108 MILITARY CENTRAL HOSPITAL
GIF-H170 | 2827553 | DAI PHUOC GENERAL CLINIC
GIF-XP170N | 2831696 | AN SINH HOSPITAL
UHI-4 | 7946492 | CAN THO PEDIATRIC HOSPITAL
GIF-H190 | 2061062 | 108 MILITARY CENTRAL HOSPITAL
GIF-H190 | 2060752 | 108 MILITARY CENTRAL HOSPITAL
GIF-EZ1500(E1) | 2306036 | 91 PHAM HUU CHI CLINIC
CF-LV1I(PAL) | 2910526 | HCM TRADITIONAL MEDICINE HOSPITAL
CF-H170L | 2339114 | GIA DINH PEOPLE HOSPITAL
GIF-H170 | 2613046 | THAI BINH PEDIATRICS HOSPITAL
GIF-H170 | 2135005 | HAI PHONG INTERNATIONAL HOSPITAL - VINH BAO
CF-H170I | 2922793 | 105 MILITARY HOSPITAL
CF-H170I | 2611051 | 105 MILITARY HOSPITAL
CF-H170L | 2823664 | BINH DINH GENERAL JOINT STOCK COMPANY
GIF-H170 | 2726150 | BINH DINH GENERAL JOINT STOCK COMPANY
GIF-H170 | 2135350 | PHUC YEN GENERAL HOSPITAL
GIF-HQ190 | 2374988 | BINH DAN HOSPITAL
GIF-EZ1500(E1) | 2408237 | DONG NAI GENERAL HOSPITAL
GIF-H170 | 2614134 | BINH DAN HOSPITAL
CF-H170I | 2435687 | VIET NAM-THUY DIEN UONG BI HOSPITAL
GIF-EZ1500(E1) | 2203810 | DOCTOR CHECK CLINIC
GIF-HQ190 | 2375129 | UNIVERSITY MEDICAL CENTER HCMC
CF-HQ190I | 2163025 | UNIVERSITY MEDICAL CENTER HCMC
UHI-4 | 7946393 | SOC TRANG PROVINCE GENERAL HOSPITAL
GIF-Q150 | 2800275 | AN GIANG REGIONAL GENERAL HOSPITAL
BF-Q190 | 2322624 | TAM ANH HOSPITAL- HCMC BRANCH
BF-1TQ170 | 2502436 | ANH KHOA MEDICAL EQUIPMENT Co., LTD
GIF-HQ190 | 2476377 | SEAGEARO LONG AN HOSPITAL
ENF-V4 | 2900480 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
GIF-EZ1500(E1) | 2001547 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
UHI-4 | 7945980 | NGUYEN DINH CHIEU HOSPITAL
GIF-XP190N | 2336542 | LOTUS CLINIC
UHI-4 | 7843794 | CAI LAY REGIONAL GENERAL HOSPITAL
UHI-4 | 7160697 | TRA VINH GENERAL HOSPITAL
GIF-HQ190 | 2374955 | UNIVERSITY MEDICAL CENTER HCMC
GIF-HQ190 | 2374941 | UNIVERSITY MEDICAL CENTER HCMC
GIF-XP170N | 2400439 | PHU THO GENERAL HOSPITAL
CF-EZ1500DI(E1) | 2101318 | TAM TRI DA NANG GENERAL HOSPITAL
CH-S200-XZ-EB | 7221522 | NGUYEN DINH CHIEU HOSPITAL
CF-EZ1500DI(E1) | 2402226 | 91 PHAM HUU CHI CLINIC
CF-H170I | 2023088 | ONCOLOGY HOSPITAL ( BRANCH 2 HCM)
CF-H170I | 2023089 | ONCOLOGY HOSPITAL ( BRANCH 2 HCM)
GIF-H170 | 2031491 | HOAN MY THU DUC INTERNATIONAL GENERAL HOSPITAL
CF-EZ1500DL(E1) | 2201706 | TAM ANH HOSPITAL- HCMC BRANCH
CF-H170I | 2922681 | THAIBINH HOSPITAL.,CO LTD
OER-AW | 2732312 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-H170 | 2338493 | HUNG VUONG GIA LAI HOSPITAL
WA2UR21A | 800202 | UNIVERSITY MEDICAL CENTER HCMC
ENF-V3 | 2612330 | PEDIATRIC HOSPITAL NO. 2
PCF-Q180AI | 2000532 | PEDIATRIC HOSPITAL NO. 2
GIF-H170 | 2032509 | THAI NGUYEN CENTRE GENERAL HOSPITAL
GIF-HQ190 | 2066604 | VINMEC PHU QUOC INTERNATIONAL HOSPITAL
CF-H170I | 2922733 | 354 MILITARY HOSPITAL
GIF-XP150N | 2502444 | CHO RAY HOSPITAL
UHI-4 | 7843589 | TAN CHAU DISTRICT GENERAL HOSPITAL - AN GIANG
MAF-TM(STERRAD NX) | 2763488 | CHO RAY HOSPITAL
SIF-Q180 | 2522609 | CHO RAY HOSPITAL
GIF-H190 | 2369525 | 87 MILITARY HOSPITAL
CF-HQ190I | 2263334 | CITY INTERNATIONAL HOSPITAL
BF-P190 | 2114397 | UNIVERSITY MEDICAL CENTER HCMC
GIF-HQ190 | 2374147 | VINMEC INTERNATIONAL HOSPITAL JSC.
GIF-HQ190 | 2613410 | VINMEC INTERNATIONAL HOSPITAL JSC.
BF-Q170 | 2502372 | UNIVERSITY MEDICAL CENTER HCMC
TJF-Q190V | 2124307 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
UHI-4 | 7843645 | THONG NHAT DONG NAI GENERAL HOSPITAL
GIF-H190 | 2061029 | VINGROUP JOINT STOCK COMPANY
CF-H190I | 2043622 | VINGROUP JOINT STOCK COMPANY
GIF-H190 | 2060737 | VINGROUP JOINT STOCK COMPANY
GIF-H190 | 2061040 | VINMEC INTERNATIONAL HOSPITAL JSC.
OFP-2 | 22020988 | CHO RAY HOSPITAL
GIF-HQ190 | 2066547 | HA NOI FRENCH HOSPITAL
GIF-H190 | 2960401 | THONG NHAT HOSPITAL
GIF-HQ190 | 2372999 | UNIVERSITY MEDICAL CENTER HCMC
A70960A | 800372 | LE VAN THINH HOSPITAL
KV-5 | 21300260 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
OTV-S400 | 7812075 | HUE CENTRAL HOSPITAL
GIF-HQ190 | 2271432 | HOAN MY SAIGON HOSPITAL
BF-1TQ170 | 2000932 | ONCOLOGY HOSPITAL ( BRANCH 2 HCM)
GIF-H170 | 2501690 | 108 MILITARY CENTRAL HOSPITAL
GIF-H190 | 2856424 | HA NOI FRENCH HOSPITAL
CF-H170I | 2822042 | CHO RAY HOSPITAL
MAJ-1430 | 7516262 | CHO RAY HOSPITAL
MAJ-1558 | 7004211 | CHO RAY HOSPITAL
MAJ-1558 | 7903405 | CHO RAY HOSPITAL
CF-H170I | 2922674 | THAI NGUYEN INTERNATIONAL HOSPITAL
CV-170 | 7098762 | DYM Medical Center Vietnam
GIF-H170 | 2031862 | VICTORIA CLINIC
ESG-400 | B003675 | UNIVERSITY MEDICAL CENTER HCMC
CF-Q150I | 2501655 | AN GIANG CENTRAL GENERAL HOSPITAL
CF-HQ190I | 2862389 | KINH BAC GENERAL HOSPITAL
CLV-190 | 7102681 | VINH LONG GENERAL HOSPITAL
CH-S700-XZ-EA | 7413843 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-Q150 | 2519420 | AN GIANG CENTRAL GENERAL HOSPITAL
CF-HQ190I | 2463633 | PHUONG CHAU INTERNATIONAL HOSPITAL
UHI-4 | 7945717 | THAP MUOI REGIONAL GENERAL HOSPITAL
CF-H170I | 2922920 | 108 MILITARY CENTRAL HOSPITAL
CF-HQ190I | 2631402 | 108 MILITARY CENTRAL HOSPITAL
MAJ-1558 | 7802495 | GIA DINH PEOPLE HOSPITAL
A22002A | 641687 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CF-EZ1500DI(E1) | 2201469 | TAM ANH HOSPITAL- HCMC BRANCH
A22002A | 734221 | THU DUC REGIONAL GENERAL HOSPITAL
BF-Q170 | 2502383 | PHAM NGOC THACH HOSPITAL
CV-190 | 7051119 | TAY NGUYEN REGIONAL GENERAL HOSPITAL
CLV-190 | 7000372 | TAY NGUYEN REGIONAL GENERAL HOSPITAL
GIF-H170 | 2134069 | THONG NHAT HOSPITAL
CF-HQ190I | 2263367 | DAI PHUOC GENERAL CLINIC
GIF-H170 | 2033400 | VUNG TAU HOSPITAL
GIF-Q150 | 2620926 | GIA LAI GENERAL HOSPITAL
CF-HQ190I | 2163164 | UNIVERSITY MEDICAL CENTER HCMC
GIF-HQ190 | 2374977 | UNIVERSITY MEDICAL CENTER HCMC
GIF-HQ190 | 2375052 | UNIVERSITY MEDICAL CENTER HCMC
GIF-EZ1500(E1) | 2102165 | DIGESTIVE & LIVER CENTER
CF-EZ1500DI(E1) | 2201387 | DA NANG HOSPITAL
GIF-EZ1500(E1) | 2203813 | DA NANG HOSPITAL
CF-EZ1500DI(E1) | 2201427 | 22-12 HOSPITAL
OTV-S200 | 7810615 | TIEN GIANG CENTRAL GENERAL HOSPITAL
GIF-EZ1500(E1) | 2305095 | FV HOSPITAL
OEV262H | 7417757 | HANH PHUC HOSPITAL
GIF-EZ1500(E1) | 2407509 | BUON MA THUOT GENERAL HOSPITAL
UHI-4 | 7843692 | PEOPLE'S HOSPITAL 115
CF-H170I | 2435694 | VAN HANH MEDICAL CENTER
CHF-V2 | 2200264 | TAM ANH HOSPITAL- HCMC BRANCH
UHI-4 | 7535804 | VINMEC PHU QUOC INTERNATIONAL HOSPITAL
Dummy-CHIP | 2101281 | DIGESTIVE & LIVER CENTER
Model | Serial No. | Enduser
BF-H190 | 2502088 | K HOSPITAL
ENF-VT2(STERRAD NX) | 2751885 | K HOSPITAL
GIF-H170 | 2401337 | K HOSPITAL
CF-H170I | 2300118 | K HOSPITAL
GIF-H170 | 2614267 | K HOSPITAL
EU-ME1 | 1121736 | E HOSPITAL
BF-UC190F | 7000581 | MILITARY HOSPITAL NO.103
BF-H190 | 2026793 | NATIONAL LUNG HOSPITAL
CF-H190I | 2621711 | K HOSPITAL ( BRANCH No.3 )
CV-170 | 7646728 | 108 MILITARY CENTRAL HOSPITAL
CYF-VA2(STERRAD NX) | 2540739 | K HOSPITAL ( BRANCH No.3 )
GIF-H170 | 2827015 | K HOSPITAL
CV-170 | 7545142 | 108 MILITARY CENTRAL HOSPITAL
CV-170 | 7545146 | 108 MILITARY CENTRAL HOSPITAL
TJF-150 | 2401510 | E HOSPITAL
OER-A | 2320996 | HA NOI FRENCH HOSPITAL
OER-AW | 2632004 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
CV-170 | 7647064 | BAC GIANG ONCOLOGY HOSPITAL
ENF-VT2 | 2040311 | ENT HOSPITAL
BF-1T150(STERRAD NX) | 2140641 | BACH MAI HOSPITAL
GIF-HQ190 | 2963456 | BAI CHAY HOSPITAL
CV-150 | 7044491 | HA DONG GENERAL HOSPITAL
CLV-180 | 7018852 | NATIONAL HOSPITAL OF PEDIATRICS
GIF-H170 | 2133781 | HAI PHONG INTERNATIONAL HOSPITAL
ENF-VT2(STERRAD NX) | 2852196 | ENT HOSPITAL
ESG-100 | B000011 | THANH NHAN HOSPITAL
CV-170 | 7658121 | E HOSPITAL
URF-V(STERRAD NX) | 2973439 | HA NOI FRENCH HOSPITAL
CV-170 | 7443025 | BACH MAI HOSPITAL
GIF-H170 | 2826996 | POLICE TRADITIONAL HOSPITAL
CF-HQ190I | 2761956 | QUANG TRI GENERAL HOSPITAL
BF-Q170 | 2500123 | VINH MEDICAL UNIVERSITY HOSPITAL
UHI-4 | 7842700 | HA DONG GENERAL HOSPITAL
CLV-180 | 7111138 | SAINT PAUL GENERAL HOSPITAL
CV-170 | 7098758 | BACH MAI HOSPITAL
GIF-HQ190 | 2961535 | THU CUC INTERNATIONAL GENERAL HOSPITAL
GIF-H170 | 2613489 | GIA BINH DISTRICT GENERAL HOSPITAL
GIF-H170 | 2726496 | VINH CITY GENERAL HOSPITAL
BF-Q170 | 2001069 | BACH MAI HOSPITAL
BF-Q170 | 2001055 | NGHE AN LUNG HOSPITAL
OER-AW | 2421393 | NATIONAL HOSPITAL OF PEDIATRICS
CH-S400-XZ-EB | 7702705 | THANH HOA PROVINCE GENERAL HOSPITAL
GIF-Q180 | 2304367 | MILITARY 4 HOSPITAL
TJF-160VR | 2403303 | 110 MILITARY HOSPITAL
TJF-150 | 2100520 | E HOSPITAL
CV-170 | 7098455 | HAI PHONG INTERNATIONAL HOSPITAL - VINH BAO
CF-H170I | 2721506 | HA TINH GENERAL HOSPITAL
CLV-180 | 7902026 | HA TINH GENERAL HOSPITAL
GIF-H170 | 2032646 | HAI PHONG INTERNATIONAL HOSPITAL
GIF-H190 | 2630170 | QUANG NINH GENERAL HOSPITAL
CF-H190I | 2631852 | QUANG NINH GENERAL HOSPITAL
GIF-H170 | 2033188 | THAI NGUYEN INTERNATIONAL HOSPITAL
CHF-V(STERRAD NX) | 2540927 | E HOSPITAL
CF-Q150I | 2301345 | BACH MAI HOSPITAL
CF-Q180AI | 2102175 | BAI CHAY HOSPITAL
GIF-H170 | 2032767 | THU CUC INTERNATIONAL HOSPITAL - LINH DAM
GIF-H170 | 2726048 | TU SON MEDICAL CENTER
CF-Q150I | 2301254 | 110 MILITARY HOSPITAL
CF-H170I | 2400151 | BACH MAI HOSPITAL
CV-180 (PAL) | 7101885 | VIET - GERMANY AMITY HOSPITAL
GIF-H170 | 2930244 | HAI DUONG GENERAL HOSPITAL
CF-H190I | 2143948 | HUNG VIET CANCER HOSPITAL
CF-H170I | 2023366 | THU CUC INTERNATIONAL CLINIC - BRANCH 2
CV-170 | 7986223 | THU CUC INTERNATIONAL CLINIC - BRANCH 2
CV-170 | 7985543 | THU CUC INTERNATIONAL CLINIC - BRANCH 2
CV-170 | 7646139 | THU CUC INTERNATIONAL CLINIC - BRANCH 2
GIF-H170 | 2930221 | SON LA GENERAL HOSPITAL
GIF-H170 | 2931168 | 108 MILITARY CENTRAL HOSPITAL
ESG-400 | B002884 | THANH NHAN HOSPITAL
GIF-H170 | 2726123 | BAC NINH PROVINCAL GENERAL HOSPITAL
EU-ME2 | 7915435 | THU CUC INTERNATIONAL GENERAL HOSPITAL
GIF-Q180 | 2625070 | QUANG NINH GENERAL HOSPITAL
BF-P150 | 2240637 | HUNG YEN GENERAL HOSPITAL
BF-P60 | 2454980 | LUNG & TUBERCULOSIS HOSPITAL OF LANG SON PROVINCE
CF-Q150I | 2800151 | NINH BINH GENERAL HOSPITAL
BF-P180(STERRAD NX) | 2343155 | NINH BINH GENERAL HOSPITAL
GIF-Q150 | 2922241 | NINH BINH GENERAL HOSPITAL
CF-Q150I | 2000441 | THAI NGUYEN UNIVERSITY OF MEDICINE HOSPITAL
ESG-400 | B006518 | P & T GENERAL HOSPITAL NO.1
GIF-H170 | 2726049 | BAC NINH PROVINCAL GENERAL HOSPITAL
CF-H170I | 2721516 | BAC NINH PROVINCAL GENERAL HOSPITAL
CV-170 | 7646344 | BAC NINH PROVINCAL GENERAL HOSPITAL
BF-P150(STERRAD NX) | 2751188 | BAC NINH PROVINCAL GENERAL HOSPITAL
CF-H170I | 2721490 | NINH BINH GENERAL HOSPITAL
GIF-H170 | 2828557 | 91 MILITARY HOSPITAL
BF-P150 | 2540972 | BAC NINH PROVINCAL GENERAL HOSPITAL
CV-180 (PAL) | 7887407 | 108 MILITARY CENTRAL HOSPITAL
GIF-H170 | 2033105 | BACH MAI HOSPITAL
OTV-S300 | 7700597 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CV-170 | 7544972 | HAI PHONG INTERNATIONAL HOSPITAL
GIF-Q150 | 2720971 | HAI PHONG INTERNATIONAL HOSPITAL
TJF-Q180V | 2619455 | SAINT PAUL GENERAL HOSPITAL
GIF-Q150 | 2002994 | HAI DUONG GENERAL HOSPITAL
GIF-H170 | 2031753 | THANH HOA PROVINCE GENERAL HOSPITAL
GIF-XP190N | 2722748 | QUANG TRI GENERAL HOSPITAL
GIF-H170 | 2930587 | NATIONAL HOSPITAL OF PEDIATRICS
CF-Q180AI | 2202420 | THANH HOA PROVINCE GENERAL HOSPITAL
CF-H170I | 2023063 | THANH HOA PROVINCE GENERAL HOSPITAL
ENF-VT2(STERRAD NX) | 2541196 | YEN BAI PROVINCIAL GENERAL HOSPITAL
CF-H170I | 2721518 | YEN BAI PROVINCIAL GENERAL HOSPITAL
CLV-S190 | 7724028 | QUANG TRI GENERAL HOSPITAL
GIF-H170 | 2612960 | HUNG YEN GENERAL HOSPITAL
GIF-H170 | 2929735 | HOP LUC GENERAL HOSPITAL
GIF-H170 | 2930325 | NGHE AN OBSTERTRICS AND PEDIATRIC HOSPITAL
GF-UCT140-AL5 | 1221568 | 108 MILITARY CENTRAL HOSPITAL
TJF-Q180V | 2404801 | DUC GIANG GENERAL HOSPITAL
TJF-Q180V | 2619551 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
GIF-XP170N | 2232176 | HAI PHONG INTERNATIONAL HOSPITAL
CV-170 | 7646140 | LUONG TAI DISTRICT GENERAL HOSPITAL
GIF-H190 | 2527126 | HA TINH GENERAL HOSPITAL
GIF-H170 | 2032651 | HAI PHONG INTERNATIONAL HOSPITAL
GIF-H170 | 2133797 | HAI PHONG INTERNATIONAL HOSPITAL
CV-170 | 7781099 | MEDLATEC GENERAL HOSPITAL
ENF-VT3 | 2900283 | 108 MILITARY CENTRAL HOSPITAL
CF-Q150I | 2930192 | CAM PHA HOSPITAL
GIF-H170 | 2930313 | NATIONAL LUNG HOSPITAL
BF-1T150 | 2651702 | BAC NINH PULMONOLOGY HOSPITAL
BF-P150(STERRAD NX) | 2651112 | BAC NINH PULMONOLOGY HOSPITAL
CF-Q150I | 2711938 | DUC GIANG GENERAL HOSPITAL
GIF-H170 | 2828971 | MEDLATEC GENERAL HOSPITAL
GIF-H170 | 2512454 | YEN BAI PROVINCIAL GENERAL HOSPITAL
CF-H190I | 2621454 | VIET NHAT PRIVATE CLINIC
OER-AW | 2732239 | VINMEC INTERNATIONAL HOSPITAL JSC.
GIF-H170 | 2031751 | THANH HOA PROVINCE GENERAL HOSPITAL
BF-P150 | 2900177 | BAC GIANG PROVINCE GENERAL HOSPITAL
CF-Q180AI | 2002021 | MILITARY HOSPITAL NO.103
GIF-H170 | 2512927 | YEN BAI PROVINCIAL GENERAL HOSPITAL
GIF-Q150 | 2001921 | NATIONAL HOSPITAL OF PEDIATRICS
CV-170 | 7009947 | BACH MAI HOSPITAL
CV-170 | 7009953 | BACH MAI HOSPITAL
ENF-V4 | 2100622 | NATIONAL HOSPITAL OF PEDIATRICS
GIF-Q150 | 2317547 | GENERAL HOSPITAL OF HOA BINH PROVINCE
GIF-H170 | 2726050 | BAC NINH PROVINCAL GENERAL HOSPITAL
CYF-V2 | 2920562 | HA TINH GENERAL HOSPITAL
BF-Q170 | 2001070 | BACH MAI HOSPITAL
GIF-H170 | 2614082 | VINH TUONG DISTRICT HOSPITAL
GIF-LV1(PAL) | 1401338 | YEN THE DISTRICT GENERAL HOSPITAL
TJF-Q180V | 2303805 | HA NOI FRENCH HOSPITAL
CF-Q150I | 2200784 | BACH MAI HOSPITAL
ESG-100 | 14308P500001 | K HOSPITAL ( BRANCH No.3 )
CLV-180 | 7110667 | VIET - GERMANY AMITY HOSPITAL
GIF-Q150 | 2800220 | THAI NGUYEN IRON AND STEEL HOSPITAL
GIF-H170 | 2300347 | GOLD STAR GENERAL HOSPITAL
CF-Q180AI | 2901629 | HA TINH GENERAL HOSPITAL
GIF-H170 | 2092189 | MILITARY HOSPITAL NO.103
GIF-Q150 | 2001872 | THAI BINH MEDICAL UNIVERSITY HOSPITAL
OTV-S7H-1N | 7527300 | THANH NHAN HOSPITAL
CLV-S45 | 7211475 | THANH NHAN HOSPITAL
CV-190 | 7620088 | SON TAY DISTRICT GENERAL HOSPITAL
CF-H170I | 2500550 | 108 MILITARY CENTRAL HOSPITAL
GIF-H180 | 2307484 | BAI CHAY HOSPITAL
BF-H190 | 2723950 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CF-H190I | 2732005 | K HOSPITAL ( BRANCH No.3 )
GIF-H190 | 2629472 | K HOSPITAL ( BRANCH No.3 )
CF-H190I | 2943414 | K HOSPITAL ( BRANCH No.3 )
CF-H190I | 2943412 | K HOSPITAL ( BRANCH No.3 )
GIF-H190 | 2742633 | K HOSPITAL ( BRANCH No.3 )
GIF-Q150 | 2721110 | HAI PHONG INTERNATIONAL HOSPITAL
CF-H170I | 2922668 | NATIONAL LUNG HOSPITAL
GIF-H190 | 2959682 | K HOSPITAL ( BRANCH No.3 )
GIF-XP170N | 2500857 | VINH MEDICAL UNIVERSITY HOSPITAL
BF-Q170 | 2700391 | LE HUU TRAC NATIONAL BURN HOSPITAL
CF-Q150I | 2401495 | MILITARY 4 HOSPITAL
BF-XP160F(STERRAD NX) | 2751617 | LE HUU TRAC NATIONAL BURN HOSPITAL
GIF-Q180 | 2625071 | QUANG NINH GENERAL HOSPITAL
GIF-XP170N | 2931776 | NATIONAL LUNG HOSPITAL
ENF-VT | 2610236 | COAL HOSPITAL
GIF-H170 | 2134150 | BACH MAI HOSPITAL
CF-Q150I | 2611919 | HAI PHONG INTERNATIONAL HOSPITAL
TJF-Q180V | 2507965 | VIET - GERMANY AMITY HOSPITAL
CF-H170I | 2721517 | BAC NINH PROVINCAL GENERAL HOSPITAL
CF-Q150I | 2501700 | E HOSPITAL
CF-Q150I | 2301390 | E HOSPITAL
TJF-150 | 2900278 | HONG NGOC GROUP - PHUC TRUONG MINH BRANCH
ENF-V4 | 2200639 | MEDLATEC GENERAL HOSPITAL
CV-150 | 7832022 | 108 MILITARY CENTRAL HOSPITAL
CV-170 | 7098607 | MILITARY HOSPITAL NO.103
GIF-H170 | 2613138 | TIEN DU MEDICAL CENTER
GIF-H170 | 2613899 | HAI DUONG UNIVERSITY OF MEDICAL TECHNOLOGY
OER-AW | 2133702 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
CF-Q150I | 2301093 | POLICE TRADITIONAL HOSPITAL
GIF-XP170N | 2400709 | JAPAN VIETNAM MEDICAL CONNECTION CO. LTD (SAKURA CLINIC )
TJF-Q180V | 2034416 | 105 MILITARY HOSPITAL
GIF-XP150N | 2612540 | SAINT PAUL GENERAL HOSPITAL
GIF-H170 | 2135441 | VINH PHUC GENERAL HOSPITAL
CF-H170I | 2124169 | VINH PHUC GENERAL HOSPITAL
CYF-VA2 | 2340619 | NINH BINH GENERAL HOSPITAL
GIF-Q150 | 2822036 | HAI PHONG INTERNATIONAL HOSPITAL
GIF-H170 | 2135440 | DAI TU DISTRICT HOSPITAL
CF-H170I | 2610936 | YEN LAC DISTRICT HOSPITAL
GIF-H170 | 2614446 | NATIONAL HOSPITAL OF ENDOCRINOLOGY ( BRANCH No.2 )
GIF-Q180 | 2514938 | 105 MILITARY HOSPITAL
UHI-4 | 7210211 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CV-170 | 7985393 | MEDLATEC GENERAL HOSPITAL
CF-H170I | 2611089 | NATIONAL HOSPITAL OF ENDOCRINOLOGY ( BRANCH No.2 )
GIF-HQ190 | 2613360 | 108 MILITARY CENTRAL HOSPITAL
CV-170 | 7098068 | VINH MEDICAL UNIVERSITY HOSPITAL
GIF-XP170N | 2031912 | K HOSPITAL ( BRANCH No.2 )
GIF-H170 | 2032515 | THAI NGUYEN CENTRE GENERAL HOSPITAL
TJF-Q180V | 2822197 | PUBLIC SECURITY MINISTRY 19/8 HOSPITAL
ESG-400 | B006130 | PHUONG BAC HOSPITAL
GIF-H170 | 2400644 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
TJF-Q180V | 2923869 | NATIONAL HOSPITAL OF PEDIATRICS
CF-H170I | 2500488 | HAI DUONG UNIVERSITY OF MEDICAL TECHNOLOGY
CV-170 | 7543810 | THANH NHAN HOSPITAL
GIF-H170 | 2930315 | NATIONAL LUNG HOSPITAL
BF-P150(STERRAD NX) | 2900159 | LANG SON PROVINCIAL GENERAL HOSPITAL
CV-170 | 7545087 | DONG HOI VIET NAM - CUBA FRENDSHIP HOSPITAL
CV-170 | 7311365 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
GIF-H190 | 2959669 | HAI PHONG CHILDREN'S HOSPITAL
GIF-Q180 | 2705920 | 108 MILITARY CENTRAL HOSPITAL
GIF-Q180 | 2203456 | DUC GIANG GENERAL HOSPITAL
TJF-Q180V | 2923384 | VINMEC HA LONG INTERNATIONAL HOSPITAL
GIF-XP180N | 2401579 | QUANG NINH OBSTETRICS AND PEDIATRICS HOSPITAL
TJF-Q180V | 2721426 | 109 MILITARY HOSPITAL
CF-Q160AI | 2210512 | QUANG NINH GENERAL HOSPITAL
GIF-HQ190 | 2858811 | VIET XO FRIENDSHIP HOSPITAL
GIF-H170 | 2929117 | NATIONAL HOSPITAL OF PEDIATRICS
WA53005A | 804616 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CF-H170I | 2123796 | BACH MAI HOSPITAL
CF-H170I | 2821908 | BACH MAI HOSPITAL
CF-H170I | 2711225 | BACH MAI HOSPITAL
GIF-H170 | 2614661 | BACH MAI HOSPITAL
CF-H170I | 2023508 | BACH MAI HOSPITAL
GIF-H170 | 2401378 | BACH MAI HOSPITAL
ENF-VT2(STERRAD NX) | 2852050 | THANH NHAN HOSPITAL
GIF-Q150 | 2721346 | BACH MAI HOSPITAL
CF-Q150I | 2900252 | BACH MAI HOSPITAL
CF-Q150I | 2601856 | BACH MAI HOSPITAL
GIF-Q150 | 2721363 | BACH MAI HOSPITAL
GIF-Q150 | 2721066 | BACH MAI HOSPITAL
GIF-Q150 | 2922411 | BACH MAI HOSPITAL
GIF-Q150 | 2821660 | BACH MAI HOSPITAL
GIF-Q150 | 2821662 | BACH MAI HOSPITAL
BF-P150 | 2700047 | THAI BINH LUNG HOSPITAL
CLV-180 | 7004130 | PUBLIC SECURITY MINISTRY 19/8 HOSPITAL
CV-180 (PAL) | 7104085 | PUBLIC SECURITY MINISTRY 19/8 HOSPITAL
PCF-PH190I | 2240411 | VINH PHUC GENERAL HOSPITAL
GIF-H170 | 2338559 | HAI PHONG INTERNATIONAL HOSPITAL
GIF-Q180 | 2525025 | VIET - GERMANY AMITY HOSPITAL
CV-180 (PAL) | 7103246 | E HOSPITAL
LTF-160 | 2600340 | HAI DUONG LUNG HOSPITAL
LTF-160(STERRAD NX) | 2051342 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
GIF-H190 | 2164739 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
GIF-H170 | 2338571 | HAI PHONG INTERNATIONAL HOSPITAL
OTV-S200 | 7355023 | MEDLATEC GENERAL HOSPITAL
CV-170 | 7545079 | VIET TIEP HOSPITAL
BF-1T150 | 2140644 | BACH MAI HOSPITAL
CF-H190I | 2144350 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
ESG-400 | 13024W130004 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CLV-180 | 7110984 | VIET - GERMANY AMITY HOSPITAL
BF-UC180F | 1822374 | NATIONAL LUNG HOSPITAL
CV-180 | 7204868 | DUC GIANG GENERAL HOSPITAL
BF-1TH190 | 2026431 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
BF-1T150(STERRAD NX) | 2052063 | K74 HOSPITAL
GIF-HQ190 | 2169025 | VINH PHUC GENERAL HOSPITAL
CYF-5A | 2901049 | HA NOI FRENCH HOSPITAL
ESG-400 | B005579 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
GIF-H190 | 2164726 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
CF-Q180AI | 2302690 | NINH BINH GENERAL HOSPITAL
GIF-H170 | 2725759 | NINH BINH GENERAL HOSPITAL
BF-P180(STERRAD NX) | 2242711 | PUBLIC SECURITY MINISTRY 19/8 HOSPITAL
CF-H170I | 2023200 | LAP THACH HEALTH CENTRE
CV-170 | 7544162 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CLH-250 | 7056701 | HEALTH PROTECTION BOARD OF HAI DUONG PROVINCE
ESG-400 | 14083W130013 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CV-170 | 7647210 | 108 MILITARY CENTRAL HOSPITAL
CF-Q150I | 2711953 | 108 MILITARY CENTRAL HOSPITAL
GIF-H190 | 2164743 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
GIF-EZ1500(E1) | 2203344 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
GIF-H170 | 2135532 | VO NHAI DISTRICT HOSPITAL
GF-UCT180 | 1121572 | HAI PHONG INTERNATIONAL HOSPITAL
OTV-S200 | 7040057 | BACH MAI HOSPITAL
OER-AW | 2832771 | VINMEC INTERNATIONAL HOSPITAL JSC.
CV-V1 | 1401393 | YEN THE DISTRICT GENERAL HOSPITAL
CV-170 | 7545656 | MEDLATEC GENERAL HOSPITAL
CF-HQ190I | 2862213 | DOLIFE INTERNATIONAL HOSPITAL
ENF-V2 | 2930334 | HA TINH GENERAL HOSPITAL
ESG-400 | B008935 | THAI NGUYEN CENTRE GENERAL HOSPITAL
OTV-S200 | 7700066 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CF-H170I | 2023679 | THAI NGUYEN INTERNATIONAL HOSPITAL
ESG-150 | B001732 | TRAN-LE TRADING AND PRODUCTION CO., LTD
GIF-HQ190 | 2168756 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
CF-H170I | 2334998 | BACH MAI HOSPITAL
BF-Q170 | 2800573 | HA NOI FRENCH HOSPITAL
BF-Q170 | 2800578 | HA NOI FRENCH HOSPITAL
CV-170 | 7645919 | YEN PHONG DISTRICT HOSPITAL
CH-S400-XZ-EB | 7702535 | THANH NHAN HOSPITAL
GIF-H170 | 2828761 | VIET NAM-THUY DIEN UONG BI HOSPITAL
OTV-S7PRO | 7478872 | VIET - GERMANY AMITY HOSPITAL
OTV-S400 | 7500356 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CV-1500 | 7022909 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
BF-H190 | 2026737 | HA NOI FRENCH HOSPITAL
GIF-Q150 | 2306580 | TUYEN QUANG POLICE HOSPITAL
CF-H170I | 2610796 | HAI DUONG UNIVERSITY OF MEDICAL TECHNOLOGY
GIF-XP170N | 2031901 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CF-H170I | 2922377 | 108 MILITARY CENTRAL HOSPITAL
GIF-H190 | 2628126 | VIET - GERMANY AMITY HOSPITAL
GIF-HQ190 | 2373723 | HUNG HA PHO NOI CLINIC
CF-HQ190I | 2363457 | HUNG HA PHO NOI CLINIC
CF-H170I | 2335485 | HANOI - DONG VAN INTERNATIONAL HOSPITAL
OTV-S400 | 7711776 | THANH HOA PROVINCE GENERAL HOSPITAL
BF-H190 | 2026727 | VINMEC INTERNATIONAL HOSPITAL JSC.
GIF-H170 | 2338537 | HUNG VUONG - PHU THO GENERAL HOSPITAL
BF-Q170 | 2900898 | VINH PHUC GENERAL HOSPITAL
GIF-Q150 | 2620622 | 108 MILITARY CENTRAL HOSPITAL
GIF-H190 | 2164715 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
GIF-H190 | 2164749 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
GIF-H190 | 2164786 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
GIF-H170 | 2828783 | 354 MILITARY HOSPITAL
GIF-H190 | 2629019 | VINMEC HAI PHONG INTERNATIONAL GENERAL HOSPITAL
CH-S200-XZ-EB | 7222180 | HUNG VIET CANCER HOSPITAL
CV-170 | 7215593 | THANH THUY DISTRICT HOSPITAL
GIF-H170 | 2135137 | DONG ANH GENERAL HOSPITAL
OTV-SC2 | 7610330 | NGHE AN POLICE HOSPITAL
OTV-S200 | 7931449 | NORTH THANG LONG HOSPITAL
GIF-EZ1500(E1) | 2305045 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
GIF-H190 | 2959681 | K HOSPITAL ( BRANCH No.3 )
CV-150 | 7044610 | E HOSPITAL
CV-170 | 7986874 | 108 MILITARY CENTRAL HOSPITAL
GIF-H190 | 2267080 | AN VIET GENERAL HOSPITAL
GIF-XP170N | 2500867 | E HOSPITAL
ENF-VT2(STERRAD NX) | 2851958 | ENT HOSPITAL
BF-H190 | 2026680 | MILITARY HOSPITAL NO.103
CV-170 | 7317615 | DONG DA GENERAL HOSPITAL
GIF-H170 | 2931169 | 108 MILITARY CENTRAL HOSPITAL
ENF-VT2(STERRAD NX) | 2751827 | ENT HOSPITAL
BF-P180(STERRAD NX) | 2142545 | SAINT PAUL GENERAL HOSPITAL
OTV-S200 | 7356265 | HAI PHONG INTERNATIONAL HOSPITAL
GIF-H170 | 2340225 | KIM DONG DISTRICT HOSPITAL
ESG-400 | B000214 | MILITARY 4 HOSPITAL
ESG-400 | B006133 | THAIBINH HOSPITAL.,CO LTD
CV-170 | 7986255 | NATIONAL LUNG HOSPITAL
GIF-H190 | 2164736 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
GIF-H190 | 2164723 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
CV-170 | 7112821 | HA NOI MEDICAL UNIVERSITY HOSPITAL
MAJ-1430 | 7111337 | HAI DUONG PEDIATRICS HOSPITAL
ESG-400 | B006136 | THAIBINH HOSPITAL.,CO LTD
OTV-S200 | 7040817 | LAC VIET FRIENDSHIP HOSPITAL
USG-410 | 7312113 | SIMACAI DISTRICT HOSPITAL
ENF-V4 | 2200656 | NATIONAL HOSPITAL OF PEDIATRICS
CF-H170I | 2234629 | HONG NGOC GROUP - PHUC TRUONG MINH BRANCH
UHI-4 | 7946490 | P & T GENERAL HOSPITAL NO.1
GIF-H170 | 2135239 | HA NOI GENERAL HOSPITAL
GIF-H170 | 2134797 | HA NOI GENERAL HOSPITAL
CV-170 | 7112834 | HA NOI GENERAL HOSPITAL
CLV-180 | 7018177 | FRIENDSHIP NGHE AN GENERAL HOSPITAL
GIF-H190 | 2164752 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
GIF-H190 | 2164755 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
GIF-Q180 | 2002309 | 360 LE HOAN CLINIC
CF-Q150I | 2800019 | 360 LE HOAN CLINIC
CF-H170I | 2023127 | HUNG VUONG - PHU THO GENERAL HOSPITAL
GIF-H170 | 2338556 | HUNG VUONG - PHU THO GENERAL HOSPITAL
GIF-H190 | 2060686 | HONG NGOC GROUP - PHUC TRUONG MINH BRANCH
CV-170 | 7316216 | BACH MAI HOSPITAL
CF-H170I | 2335001 | BACH MAI HOSPITAL
GIF-H170 | 2338353 | BACH MAI HOSPITAL
GIF-H170 | 2931167 | 108 MILITARY CENTRAL HOSPITAL
UHI-4 | 7160074 | AGRICULTURE HOSPITAL
CV-180 | 7103225 | HAI DUONG PEDIATRICS HOSPITAL
CF-H170I | 2023312 | THAI NGUYEN CENTRE GENERAL HOSPITAL
TJF-Q180V | 2721718 | VINMEC HAI PHONG INTERNATIONAL GENERAL HOSPITAL
ENF-V4 | 2200635 | MEDLATEC GENERAL HOSPITAL
CV-170 | 7986168 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
BF-XP160F(STERRAD NX) | 2651581 | QUANG NINH OBSTETRICS AND PEDIATRICS HOSPITAL
CF-EZ1500DI(E1) | 2301957 | LAC VIET FRIENDSHIP HOSPITAL
CV-170 | 7781890 | NGOC LAC REGIONAL GENERAL HOSPITAL
CF-H170I | 2335011 | BACH MAI HOSPITAL
CF-H170I | 2334916 | BACH MAI HOSPITAL
CV-170 | 7431734 | BACH MAI HOSPITAL
CV-170 | 7545127 | E HOSPITAL
CV-170 | 7780273 | HAI PHONG INTERNATIONAL HOSPITAL
GIF-H170 | 2338541 | HUNG VUONG - PHU THO GENERAL HOSPITAL
GIF-H170 | 2338557 | THANH HOA PROVINCE GENERAL HOSPITAL
CF-H170I | 2234624 | HUNG VUONG - PHU THO GENERAL HOSPITAL
UHI-4 | 7740967 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-H170 | 2031859 | TOMEC HIGH TECHNOLOGY MEDICAL CENTER
ENF-VT2(STERRAD NX) | 2340861 | HA NOI MEDICAL UNIVERSITY HOSPITAL
UHI-4 | 7322997 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-H170 | 2501905 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
GIF-Q150 | 2104537 | NINH BINH GENERAL HOSPITAL
GIF-H170 | 2340231 | NINH BINH GENERAL HOSPITAL
CV-170 | 7545305 | THAI BINH MEDICAL UNIVERSITY HOSPITAL
CF-LV1I(PAL) | 2910667 | COAL HOSPITAL
CV-190 | 7629866 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CF-H170I | 2335022 | BACH MAI HOSPITAL
GIF-H170 | 2338424 | BACH MAI HOSPITAL
GIF-H190 | 2753271 | VINMEC HAI PHONG INTERNATIONAL GENERAL HOSPITAL
UHI-4 | 7740963 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
ESG-400 | B001520 | HA NOI FRENCH HOSPITAL
CF-HQ190I | 2062809 | TOMEC HIGH TECHNOLOGY MEDICAL CENTER
CH-S700-XZ-EA | 7301754 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
UHI-4 | 7740989 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
BF-P150(STERRAD NX) | 2751161 | HANOI LUNG HOSPITAL
BF-1T150(STERRAD NX) | 2240902 | HANOI LUNG HOSPITAL
GIF-XP170N | 2332312 | PHUONG BAC HOSPITAL
GIF-H190 | 2267483 | VIET - GERMANY AMITY HOSPITAL
CLV-190 | 7669086 | SAINT PAUL GENERAL HOSPITAL
MAF-TM(STERRAD NX) | 2452464 | K74 HOSPITAL
CV-180 (PAL) | 7206093 | K74 HOSPITAL
CV-180 (PAL) | 7889553 | K74 HOSPITAL
CLV-180 | 7215094 | K74 HOSPITAL
GIF-H170 | 2338425 | BACH MAI HOSPITAL
GIF-H170 | 2033025 | TTH GENERAL HOSPITAL HA TINH
GIF-Q150 | 2002991 | HA DONG GENERAL HOSPITAL
GIF-Q150 | 2205622 | HA DONG GENERAL HOSPITAL
CF-Q150I | 2200976 | HA DONG GENERAL HOSPITAL
OER-AW | 2033171 | MILITARY HOSPITAL NO.103
UHI-4 | 7638148 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CV-170 | 7009044 | VIET TIEP HOSPITAL
EL 27 | 99230143 | VIET - GERMANY AMITY HOSPITAL
GIF-H170 | 2033016 | THAI NGUYEN INTERNATIONAL HOSPITAL
GIF-H170 | 2715699 | HONG NGOC SAVICO LONG BIEN INTERNATIONAL CLINIC
UHI-4 | 7741794 | NGHIA LO GENERAL HOSPITAL
OTV-S700 | 7302151 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
OTV-S700 | 7301625 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
BF-P150 | 2600025 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
GIF-XP170N | 2332297 | K HOSPITAL
CV-170 | 7985396 | THAO NGUYEN GENERAL HOSPITAL
GIF-H170 | 2338495 | BACH MAI HOSPITAL
GIF-H170 | 2338491 | BACH MAI HOSPITAL
GIF-H170 | 2338517 | BACH MAI HOSPITAL
CF-H170I | 2335024 | BACH MAI HOSPITAL
A22002A | 727904 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CF-H170I | 2821936 | MEDLATEC GENERAL HOSPITAL
CHF-V2 | 2300490 | MILITARY HOSPITAL NO.103
CV-170 | 7884164 | MILITARY INSTITUTE OF RADIATION MEDICINE AND CANCER
BF-P180 | 2801471 | K74 HOSPITAL
CF-EZ1500DI(E1) | 2201484 | TRAN-LE TRADING AND PRODUCTION CO., LTD
GIF-HQ190 | 2373786 | BRANCH OF VINMEC INTERNATIONAL HOSPITAL JOINT STOCK COMPANY - VINMEC TIMES INTERNATIONAL HOSPITAL
GIF-H170 | 2338705 | NATIONAL LUNG HOSPITAL
CF-HQ190I | 2363460 | BRANCH OF VINMEC INTERNATIONAL HOSPITAL JOINT STOCK COMPANY - VINMEC TIMES INTERNATIONAL HOSPITAL
UHI-4 | 7944893 | P & T GENERAL HOSPITAL NO.1
ESG-150 | B002149 | AN DUONG DISTRICT GENERAL HOSPITAL
GIF-H170 | 2033205 | TTH GENERAL HOSPITAL HA TINH
CF-Q150I | 2601849 | SAINT PAUL GENERAL HOSPITAL
GIF-H170 | 2614077 | TAM DAO GENERAL HOSPITAL
CV-170 | 7986224 | HAI PHONG INTERNATIONAL HOSPITAL
UHI-4 | 7638135 | SON TAY DISTRICT GENERAL HOSPITAL
CF-HQ190I | 2062996 | HUNG HA GENERAL HOSPITAL
UHI-4 | 7945537 | THAIBINH HOSPITAL.,CO LTD
OER-AW | 2434597 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CF-H170I | 2335257 | MILITARY HOSPITAL NO.103
GIF-Q180 | 2102869 | VIET - GERMANY AMITY HOSPITAL
CF-H170I | 2721581 | ME LINH GENERAL HOSPITAL
CV-180 | 7100594 | NATIONAL HOSPITAL OF PEDIATRICS
UHI-4 | 7638121 | SON TAY DISTRICT GENERAL HOSPITAL
ESG-400 | 14077W130013 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GIF-XP180N | 2100714 | SAINT PAUL GENERAL HOSPITAL
CF-H170I | 2124155 | HA NOI GENERAL HOSPITAL
UHI-4 | 7945543 | THAIBINH HOSPITAL.,CO LTD
CF-HQ190I | 2862066 | VINMEC HAI PHONG INTERNATIONAL GENERAL HOSPITAL
A70960A | 801219 | LAO CAI OBSTERTRICS AND PEDIATRICS HOSPITAL
OER-AW | 2531653 | VINMEC INTERNATIONAL HOSPITAL JSC.
CF-EZ1500DI(E1) | 2302038 | BAI CHAY HOSPITAL
CV-170 | 7545656 | MEDLATEC CLINIC (TAY HO Branch)
CF-H170I | 2610742 | THU CUC INTERNATIONAL CLINIC - BRANCH 2
UHI-4 | 7649812 | HA NOI FRENCH HOSPITAL
GIF-H170 | 2714846 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
ESG-150 | B000477 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CV-1500 | 7352024 | LAC VIET FRIENDSHIP HOSPITAL
BF-P190 | 2415535 | 108 MILITARY CENTRAL HOSPITAL
UHI-4 | 7649815 | HA NOI FRENCH HOSPITAL
CF-H170I | 2922467 | MEDLATEC CLINIC (CAU GIAY branch)
CF-H170I | 2822045 | 354 MILITARY HOSPITAL
OTV-S400 | 7711528 | THANH NHAN HOSPITAL
CF-H190I | 2144356 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
MAJ-1558 | 7711702 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
MAJ-1430 | 7727741 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
GIF-HQ190 | 2271429 | TMC CLINIC ( T-MATSUOKA)
UHI-4 | 7946125 | 105 MILITARY HOSPITAL
CLV-190 | 7000363 | SAINT PAUL GENERAL HOSPITAL
GIF-H190 | 2061902 | K HOSPITAL
GF-UE190 | 7220529 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GF-UCT180 | 1921520 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
ESG-150 | B002475 | H Plus International Medical Center
CV-170 | 7318391 | TAM DUC CAU QUAN GENERAL HOSPITAL
CHF-V2 | 2300578 | MILITARY HOSPITAL NO.103
CH-S200-XZ-EA | 7911473 | BACH MAI HOSPITAL
ESG-400 | B000525 | SAINT PAUL GENERAL HOSPITAL
ESG-400 | B008923 | AGRICULTURE HOSPITAL
GIF-H170 | 2338719 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CF-H170I | 2234716 | THANH NHAN HOSPITAL
CLV-180 | 7315590 | MINH HOA QUANG BINH HOSPITAL
CV-180 (PAL) | 7306374 | MINH HOA QUANG BINH HOSPITAL
GIF-EZ1500(E1) | 2203320 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
UHI-4 | 7843643 | THANH NHAN HOSPITAL
UHI-4 | 7741229 | THANH NHAN HOSPITAL
CV-170 | 7986876 | 108 MILITARY CENTRAL HOSPITAL
CLV-190 | 7669092 | SAINT PAUL GENERAL HOSPITAL
CF-H170I | 2335034 | HAI PHONG INTERNATIONAL HOSPITAL
TJF-Q180V | 2721245 | QUANG TRI GENERAL HOSPITAL
CF-H170I | 2510660 | HUNG YEN GENERAL HOSPITAL
ENF-V4 | 2200639 | MEDLATEC CLINIC (THANH XUAN branch)
BF-Q170 | 2001021 | NATIONAL LUNG HOSPITAL
BF-Q170 | 2402314 | NATIONAL LUNG HOSPITAL
CF-H190I | 2144349 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
GIF-H190 | 2061914 | HONG NGOC GENERAL HOSPITAL
CV-170 | 7646784 | HAI DUONG MEDICAL TECHNICAL UNIVERSITY
CV-170 | 7545698 | NGHE AN OBSTERTRICS AND PEDIATRIC HOSPITAL
UHI-4 | 7844202 | NGHE AN OBSTERTRICS AND PEDIATRIC HOSPITAL
BF-H190 | 2026883 | SAINT PAUL GENERAL HOSPITAL
BF-H190 | 2026800 | 105 MILITARY HOSPITAL
CV-1500 | 7249389 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
WM-NP2 | 21211399 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
GF-UCT180 | 1121578 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
CV-170 | 7098530 | HA TINH GENERAL HOSPITAL
CF-H170I | 2023671 | PHO NOI GENERAL HOSPITAL
GIF-H170 | 2340573 | DONG DA GENERAL HOSPITAL
BF-Q170 | 2402305 | NATIONAL LUNG HOSPITAL
ESG-400 | B002289 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
BF-Q170 | 2201828 | NINH BINH GENERAL HOSPITAL
CF-HQ190I | 2962639 | P & T GENERAL HOSPITAL NO.1
GIF-HQ190 | 2963818 | P & T GENERAL HOSPITAL NO.1
GIF-H170 | 2236964 | THANH NHAN HOSPITAL
CV-170 | 7316213 | BACH MAI HOSPITAL
LF-TP(STERRAD) | 2026799 | LUNG & TUBERCULOSIS HOSPITAL OF LANG SON PROVINCE
GIF-H170 | 2441001 | GENERAL HOSPITAL OF HOA BINH PROVINCE
ESG-400 | B008560 | DAN PHUONG DISTRICT GENERAL HOSPITAL
CLV-S40PRO | 7528627 | BAC GIANG PROVINCE GENERAL HOSPITAL
CV-170 | 7110886 | TTH GENERAL HOSPITAL HA TINH
BF-P150(STERRAD NX) | 2040236 | HA TRUNG GENERAL HOSPITAL
UHI-4 | 7475674 | NAM DAN HOSPITAL
UHI-4 | 7841967 | THANH HOA PROVINCE GENERAL HOSPITAL
GIF-Q150 | 2721093 | NATIONAL INSTITUTE OF HEMATOLOGY AND BLOOD TRANSFUSION
UHI-4 | 7946217 | Airforce hospital
GIF-H190 | 2165088 | PUBLIC SECURITY MINISTRY 19/8 HOSPITAL
GIF-H190 | 2369679 | YEN LAC DISTRICT HOSPITAL
BF-Q170 | 2500115 | VIET TIEP HOSPITAL
GIF-H190 | 2470314 | PHENIKAA UNIVERSITY HOSPITAL
GIF-H170 | 2032956 | HOAN MY INTERNATIONAL HOSPITAL Co.,JSC
CF-H190I | 2144360 | PUBLIC SECURITY MINISTRY 19/8 HOSPITAL
GIF-H170 | 2614070 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
GIF-H170 | 2440716 | GIA LAM GENERAL HOSPITAL
BF-Q170 | 2302066 | PHENIKAA UNIVERSITY HOSPITAL
GIF-H170 | 2032456 | TOAN TAM INTERNATIONAL GENERAL CLINIC
CLV-190 | 7411142 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
BF-P190 | 2812450 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
CLV-190 | 7410283 | 108 MILITARY CENTRAL HOSPITAL
UHI-4 | 7648806 | SAINT PAUL GENERAL HOSPITAL
BF-H190 | 2026739 | K HOSPITAL
CLV-180 | 7018628 | SAINT PAUL GENERAL HOSPITAL
UHI-4 | 7945532 | BACH MAI HOSPITAL
CV-190 | 7161131 | QUANG NINH GENERAL HOSPITAL
GIF-H170 | 2440959 | VIET NAM-THUY DIEN UONG BI HOSPITAL
BF-Q170 | 2800537 | HANOI LUNG HOSPITAL
CLL-S700 | 7301584 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD
CF-H170I | 2435722 | NINH BINH GENERAL HOSPITAL
UHI-4 | 7160440 | CHUONG MY DISTRICT GENERAL HOSPITAL
TJF-Q180V | 2721421 | 354 MILITARY HOSPITAL
GIF-HQ190 | 2746465 | VINMEC INTERNATIONAL HOSPITAL JSC.
GIF-H170 | 2133588 | THAI NGUYEN CENTRE GENERAL HOSPITAL
CF-Q150I | 2000410 | 105 MILITARY HOSPITAL
GIF-HQ190 | 2373795 | TMC CLINIC ( T-MATSUOKA)
ESG-400 | B006425 | Airforce hospital
MAJ-2380 | 7411792 | 108 MILITARY CENTRAL HOSPITAL
CF-H170I | 2335295 | 354 MILITARY HOSPITAL
GIF-H170 | 2929429 | MEDLATEC CLINIC (TAY HO Branch)
BF-1T150(STERRAD NX) | 2040274 | PUBLIC SECURITY MINISTRY 19/8 HOSPITAL
GF-UE160-AL5 | 1212294 | 108 MILITARY CENTRAL HOSPITAL
UHI-4 | 7158854 | HA NOI - CAO BANG CLINIC
CF-H190I | 2445369 | OLYMPUS MEDICAL SYSTEMS VIETNAM CO., LTD
CF-H170I | 2721542 | YEN BAI PROVINCIAL GENERAL HOSPITAL
BF-P190 | 2415578 | 108 MILITARY CENTRAL HOSPITAL
CV-170 | 7543685 | 108 MILITARY CENTRAL HOSPITAL
GIF-H170 | 2440932 | VINH CITY GENERAL HOSPITAL
GIF-EZ1500(E1) | 2304566 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
CV-170 | 7544968 | HAI DUONG MEDICAL TECHNICAL UNIVERSITY
CF-H170I | 2335367 | MILITARY 4 HOSPITAL
CF-EZ1500DI(E1) | 2201474 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
GIF-Q180 | 2002208 | 105 MILITARY HOSPITAL
GIF-H170 | 2613887 | BA CHE MEDICAL CENTRE
GIF-H170 | 2133751 | TTH GENERAL HOSPITAL HA TINH
GIF-H170 | 2133808 | TTH GENERAL HOSPITAL HA TINH
CF-EZ1500DI(E1) | 2201454 | TAM ANH GENERAL HOSPITAL JOINT STOCK COMPANY
CF-Q150I | 2301334 | 108 MILITARY CENTRAL HOSPITAL
CF-Q150I | 2600013 | 108 MILITARY CENTRAL HOSPITAL
GIF-Q150 | 2620621 | 108 MILITARY CENTRAL HOSPITAL
CV-170 | 7419323 | DONG HUNG DISTRICT GENERAL HOSPITAL
GIF-H170 | 2236611 | HAI PHONG INTERNATIONAL HOSPITAL
CF-H170I | 2435557 | DONG DA GENERAL HOSPITAL
UHI-4 | 7946762 | HUNG CUONG GENERAL HOSPITAL
OER-AW | 2434830 | VINMEC TIMES CITY INTERNATIONAL HOSPITAL
OER-AW | 2033173 | TMC CLINIC ( T-MATSUOKA)
CH-S400-XZ-EB | 7130935 | THANH HOA PROVINCE GENERAL HOSPITAL
OER-AW | 2534894 | FAMILY MEDICAL PRACTICE HANOI
GIF-H170 | 2613712 | BAC HA DISTRICT HOSPITAL
CF-H190I | 2621671 | SAINT PAUL GENERAL HOSPITAL
CF-EZ1500DI(E1) | 2402420 | VIET NAM-THUY DIEN UONG BI HOSPITAL
GIF-H190 | 2470402 | 108 MILITARY CENTRAL HOSPITAL
CF-H170I | 2123728 | THAI NGUYEN INTERNATIONAL HOSPITAL
OER-AW | 2632009 | SAINT PAUL GENERAL HOSPITAL
OER-AW | 2632010 | SAINT PAUL GENERAL HOSPITAL
BF-1T150(STERRAD NX) | 2040273 | PUBLIC SECURITY MINISTRY 19/8 HOSPITAL
OTV-S200 | 7810469 | NATIONAL TRANSPORT HOSPITAL
BF-1TQ170 | 2000998 | VIET NAM-THUY DIEN UONG BI HOSPITAL
CLV-190 | 7878453 | BAI CHAY HOSPITAL
PCF-H190DL | 2147636 | BAI CHAY HOSPITAL
GIF-H170 | 2340587 | VIET XO FRIENDSHIP HOSPITAL
GIF-H170 | 2340249 | 354 MILITARY HOSPITAL
GIF-HQ190 | 2374999 | CAM PHA HOSPITAL
UHI-4 | 7842674 | NATIONAL TRANSPORT HOSPITAL
Dummy-CHIP | 2135352 | PUBLIC SECURITY MINISTRY 19/8 HOSPITAL
0 | 0 | 0
KV-6 | 1812816 | OVN
WMP-19K-MXM | 20100115 | OVN
CF-H170I | 2821768 | MINISTRY OF PUBLIC SECURITY 199 HOSPITAL
UHI-4 | 7364293 | PHUC HUNG GENERAL HOSPITAL

## GIF-190,1200,1500
BÁO CÁO GIÁM ĐỊNH KỸ THUẬT
TECHNICAL REPORT
SỐ/ No.: |  |  |  |  | NGÀY RA BÁO CÁO/ Issued Date:
KÍNH GỞI/ Report to: |  |  | Quý Khách Hàng/ Valued Customer
BỆNH VIỆN/ Name Of User:
KIỂU MÁY/ Name Of Model: |  |  |  |  | SỐ MÁY/ Serial No:
HƯ HỎNG TỪ KHÁCH HÀNG/ Customer Complaints:
1. 2.
ĐÁNH GIÁ TỪ OLYMPUS/ Olympus findings
Bộ Phận/ Section |  | Đánh Giá Hư Hỏng/ Confirmed Defects |  |  |  |  | Hình Ảnh/ Photos
Bộ phận điều khiển Control Unit |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Yêu cầu sửa chữa/ Repair required
Dây dẫn chung Universal Cord |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Khuyến nghị sửa chữa/ Repair recommended
GHI CHÚ/ Remark
NGUYÊN NHÂN HƯ HỎNG/ Reason Of Damages
- Các hư hỏng do hao mòn trong quá trình sử dụng/ The damages due to worn and tear during use.
GIẢI PHÁP/ Solutions
- Thay thế linh kiện như danh mục chuẩn đính kèm. Tuy nhiên nếu khách hàng không đồng ý sửa chữa lớn, chúng tôi chỉ thay thế linh kiện như danh mục khuyến nghị kèm theo. Chúng tôi chỉ bảo hành cho các linh kiện thay thế này.
Replace new spare parts as standard attachment list. However, if the customer does not agree to Major repair, we will replace the new spare parts as recommended attachment list. We only warrant replaced parts.
HÀNH ĐỘNG NGĂN NGỪA/ Recommended Actions To Prevent Sử dụng, bảo quản và vệ sinh tiệt trùng thiết bị theo đúng hướng dẫn của Olympus. Use, maintain and reprocess the equipment following Olympus’s instructions.
- Sử dụng, bảo quản và vệ sinh tiệt trùng thiết bị theo đúng hướng dẫn của Olympus. - Use, maintain and reprocess the equipment following Olympus’s instructions.
THÔNG TIN SỬA CHỮA/ Repair Information
| SỬA CHỮA THEO NHU CẦU On-demand Repair |  |  | TRONG BẢO HÀNH Under Warranty |  |  |  | BẢO HÀNH MUA MỚI Product Warranty
BẢO HÀNH SỬA CHỮA Repair Warranty
KHÔNG ĐƯỢC BẢO HÀNH No Warranty
HẾT BẢO HÀNH Out Of Warranty
| HỢP ĐỒNG BẢO TRÌ Service Contract |  |  | THUỘC ĐIỀU KHOẢN HỢP ĐỒNG No-charge
NGOÀI ĐIỀU KHOẢN HỢP ĐỒNG Charge
THUỘC MỘT PHẦN ĐIỀU KHOẢN HỢP ĐỒNG Particular Charge
ĐẠI DIỆN CÔNG TY OLYMPUS | [5 empty] | ĐẠI DIỆN KHÁCH HÀNG
Representative of Olympus Company | [5 empty] | Representative of Customer
NGƯỜI KIỂM TRA/ Inspected By
South Region Technical Service Engineer
NGƯỜI DUYỆT/ Approved By
South Region Technical Service Assistant Manager
OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD. (OVN)
Office and Service Center in Hanoi: Unit 305- 307, 3rd Floor, V Tower, 649 Kim Ma Street, Giang Vo Ward, HaNoi City, Vietnam - Tel: (84-24) 3766 5831/32/34
Office In HoChiMinh City: 7th Floor, Viettel Building, 285 Cach Mang Thang Tam, Hoa Hung Ward, HoChiMinh City, Vietnam - Tel: (84-28) 7301 9919
Service Center In HoChiMinh City: 2nd Floor, 86 Tan Da, Cho Lon Ward, HoChiMinh City, Vietnam - Tel: (84-28) 7301 9919
Scope Structure
Endoscope Connector
Universal Cord
Control Section
Insertion Tube
Bending Section
Distal End

## CF&PCF-170
BÁO CÁO GIÁM ĐỊNH KỸ THUẬT
TECHNICAL REPORT
SỐ/ No.: |  |  |  |  | NGÀY RA BÁO CÁO/ Issued Date:
KÍNH GỞI/ Report to: |  |  | Quý Khách Hàng/ Valued Customer
BỆNH VIỆN/ Name Of User:
KIỂU MÁY/ Name Of Model: |  |  |  |  | SỐ MÁY/ Serial No:
HƯ HỎNG TỪ KHÁCH HÀNG/ Customer Complaints:
1. 2.
ĐÁNH GIÁ TỪ OLYMPUS/ Olympus findings
Bộ Phận/ Section |  | Đánh Giá Hư Hỏng/ Confirmed Defects |  |  |  |  | Hình Ảnh/ Photos
Bộ phận uốn cong Bending Section |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Yêu cầu sửa chữa/ Repair required
Bộ phận điều khiển Control Unit |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Yêu cầu sửa chữa/ Repair required
Dây dẫn chung Universal Cord |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Khuyến nghị sửa chữa/ Repair recommended
GHI CHÚ/ Remark
NGUYÊN NHÂN HƯ HỎNG/ Reason Of Damages
- Các hư hỏng do hao mòn trong quá trình sử dụng/ The damages due to worn and tear during use.
GIẢI PHÁP/ Solutions
- Thay thế linh kiện như danh mục chuẩn đính kèm. Tuy nhiên nếu khách hàng không đồng ý sửa chữa lớn, chúng tôi chỉ thay thế linh kiện như danh mục khuyến nghị kèm theo. Chúng tôi chỉ bảo hành cho các linh kiện thay thế này.
Replace new spare parts as standard attachment list. However, if the customer does not agree to Major repair, we will replace the new spare parts as recommended attachment list. We only warrant replaced parts.
HÀNH ĐỘNG NGĂN NGỪA/ Recommended Actions To Prevent Sử dụng, bảo quản và vệ sinh tiệt trùng thiết bị theo đúng hướng dẫn của Olympus. Use, maintain and reprocess the equipment following Olympus’s instructions.
Sử dụng, bảo quản và vệ sinh tiệt trùng thiết bị theo đúng hướng dẫn của Olympus. Use, maintain and reprocess the equipment following Olympus’s instructions.
THÔNG TIN SỬA CHỮA/ Repair Information
| SỬA CHỮA THEO NHU CẦU On-demand Repair |  |  | TRONG BẢO HÀNH Under Warranty |  |  |  | BẢO HÀNH MUA MỚI Product Warranty
BẢO HÀNH SỬA CHỮA Repair Warranty
KHÔNG ĐƯỢC BẢO HÀNH No Warranty
HẾT BẢO HÀNH Out Of Warranty
| HỢP ĐỒNG BẢO TRÌ Service Contract |  |  | THUỘC ĐIỀU KHOẢN HỢP ĐỒNG No-charge
NGOÀI ĐIỀU KHOẢN HỢP ĐỒNG Charge
THUỘC MỘT PHẦN ĐIỀU KHOẢN HỢP ĐỒNG Particular Charge
ĐẠI DIỆN CÔNG TY OLYMPUS | [5 empty] | ĐẠI DIỆN KHÁCH HÀNG
Representative of Olympus Company | [5 empty] | Representative of Customer
NGƯỜI KIỂM TRA/ Inspected By
South Region Technical Service Engineer
NGƯỜI DUYỆT/ Approved By
South Region Technical Service Assistant Manager
OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD. (OVN)
Office and Service Center in Hanoi: Unit 305- 307, 3rd Floor, V Tower, 649 Kim Ma Street, Giang Vo Ward, HaNoi City, Vietnam - Tel: (84-24) 3766 5831/32/34
Office In HoChiMinh City: 7th Floor, Viettel Building, 285 Cach Mang Thang Tam, Hoa Hung Ward, HoChiMinh City, Vietnam - Tel: (84-28) 7301 9919
Service Center In HoChiMinh City: 2nd Floor, 86 Tan Da, Cho Lon Ward, HoChiMinh City, Vietnam - Tel: (84-28) 7301 9919
Scope Structure
LG Connector
Universal Cord
Control Section
Insertion Tube
Bending Section
Distal End
Video Connector
Video Cable

## GIF-170
BÁO CÁO GIÁM ĐỊNH KỸ THUẬT
TECHNICAL REPORT
SỐ/ No.: |  |  |  |  | NGÀY RA BÁO CÁO/ Issued Date:
KÍNH GỞI/ Report to: |  |  | Quý Khách Hàng/ Valued Customer
BỆNH VIỆN/ Name Of User:
KIỂU MÁY/ Name Of Model: |  |  |  |  | SỐ MÁY/ Serial No:
HƯ HỎNG TỪ KHÁCH HÀNG/ Customer Complaints:
1. 2.
ĐÁNH GIÁ TỪ OLYMPUS/ Olympus findings
Bộ Phận/ Section |  | Đánh Giá Hư Hỏng/ Confirmed Defects |  |  |  |  | Hình Ảnh/ Photos
Bộ phận uốn cong Bending Section |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Yêu cầu sửa chữa/ Repair required
Bộ phận điều khiển Control Unit |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Yêu cầu sửa chữa/ Repair required
Dây dẫn chung Universal Cord |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Khuyến nghị sửa chữa/ Repair recommended
GHI CHÚ/ Remark
NGUYÊN NHÂN HƯ HỎNG/ Reason Of Damages
- Các hư hỏng do hao mòn trong quá trình sử dụng/ The damages due to worn and tear during use.
GIẢI PHÁP/ Solutions
- Thay thế linh kiện như danh mục chuẩn đính kèm. Tuy nhiên nếu khách hàng không đồng ý sửa chữa lớn, chúng tôi chỉ thay thế linh kiện như danh mục khuyến nghị kèm theo. Chúng tôi chỉ bảo hành cho các linh kiện thay thế này.
Replace new spare parts as standard attachment list. However, if the customer does not agree to Major repair, we will replace the new spare parts as recommended attachment list. We only warrant replaced parts.
HÀNH ĐỘNG NGĂN NGỪA/ Recommended Actions To Prevent Sử dụng, bảo quản và vệ sinh tiệt trùng thiết bị theo đúng hướng dẫn của Olympus. Use, maintain and reprocess the equipment following Olympus’s instructions.
Sử dụng, bảo quản và vệ sinh tiệt trùng thiết bị theo đúng hướng dẫn của Olympus. Use, maintain and reprocess the equipment following Olympus’s instructions.
THÔNG TIN SỬA CHỮA/ Repair Information
| SỬA CHỮA THEO NHU CẦU On-demand Repair |  |  | TRONG BẢO HÀNH Under Warranty |  |  |  | BẢO HÀNH MUA MỚI Product Warranty
BẢO HÀNH SỬA CHỮA Repair Warranty
KHÔNG ĐƯỢC BẢO HÀNH No Warranty
HẾT BẢO HÀNH Out Of Warranty
| HỢP ĐỒNG BẢO TRÌ Service Contract |  |  | THUỘC ĐIỀU KHOẢN HỢP ĐỒNG No-charge
NGOÀI ĐIỀU KHOẢN HỢP ĐỒNG Charge
THUỘC MỘT PHẦN ĐIỀU KHOẢN HỢP ĐỒNG Particular Charge
ĐẠI DIỆN CÔNG TY OLYMPUS | [5 empty] | ĐẠI DIỆN KHÁCH HÀNG
Representative of Olympus Company | [5 empty] | Representative of Customer
NGƯỜI KIỂM TRA/ Inspected By
South Region Technical Service Engineer
NGƯỜI DUYỆT/ Approved By
South Region Technical Service Assistant Manager
OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD. (OVN)
Office and Service Center in Hanoi: Unit 305- 307, 3rd Floor, V Tower, 649 Kim Ma Street, Giang Vo Ward, HaNoi City, Vietnam - Tel: (84-24) 3766 5831/32/34
Office In HoChiMinh City: 7th Floor, Viettel Building, 285 Cach Mang Thang Tam, Hoa Hung Ward, HoChiMinh City, Vietnam - Tel: (84-28) 7301 9919
Service Center In HoChiMinh City: 2nd Floor, 86 Tan Da, Cho Lon Ward, HoChiMinh City, Vietnam - Tel: (84-28) 7301 9919
Scope Structure
LG Connector
Universal Cord
Control Section
Insertion Tube
Bending Section
Distal End
Video Connector
Video Cable

## CF&PCF-180
BÁO CÁO GIÁM ĐỊNH KỸ THUẬT
TECHNICAL REPORT
SỐ/ No.: |  |  |  |  | NGÀY RA BÁO CÁO/ Issued Date:
KÍNH GỞI/ Report to: |  |  | Quý Khách Hàng/ Valued Customer
BỆNH VIỆN/ Name Of User:
KIỂU MÁY/ Name Of Model: |  |  |  |  | SỐ MÁY/ Serial No:
HƯ HỎNG TỪ KHÁCH HÀNG/ Customer Complaints:
1. 2.
ĐÁNH GIÁ TỪ OLYMPUS/ Olympus findings
Bộ Phận/ Section |  | Đánh Giá Hư Hỏng/ Confirmed Defects |  |  |  |  | Hình Ảnh/ Photos
Bộ phận uốn cong Bending Section |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Yêu cầu sửa chữa/ Repair required
Bộ phận điều khiển Control Unit |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Yêu cầu sửa chữa/ Repair required
Dây dẫn chung Universal Cord |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Khuyến nghị sửa chữa/ Repair recommended
GHI CHÚ/ Remark
NGUYÊN NHÂN HƯ HỎNG/ Reason Of Damages
- Các hư hỏng do hao mòn trong quá trình sử dụng/ The damages due to worn and tear during use.
GIẢI PHÁP/ Solutions
- Thay thế linh kiện như danh mục chuẩn đính kèm. Tuy nhiên nếu khách hàng không đồng ý sửa chữa lớn, chúng tôi chỉ thay thế linh kiện như danh mục khuyến nghị kèm theo. Chúng tôi chỉ bảo hành cho các linh kiện thay thế này.
Replace new spare parts as standard attachment list. However, if the customer does not agree to Major repair, we will replace the new spare parts as recommended attachment list. We only warrant replaced parts.
HÀNH ĐỘNG NGĂN NGỪA/ Recommended Actions To Prevent Sử dụng, bảo quản và vệ sinh tiệt trùng thiết bị theo đúng hướng dẫn của Olympus. Use, maintain and reprocess the equipment following Olympus’s instructions.
Sử dụng, bảo quản và vệ sinh tiệt trùng thiết bị theo đúng hướng dẫn của Olympus. Use, maintain and reprocess the equipment following Olympus’s instructions.
THÔNG TIN SỬA CHỮA/ Repair Information
| SỬA CHỮA THEO NHU CẦU On-demand Repair |  |  | TRONG BẢO HÀNH Under Warranty |  |  |  | BẢO HÀNH MUA MỚI Product Warranty
BẢO HÀNH SỬA CHỮA Repair Warranty
KHÔNG ĐƯỢC BẢO HÀNH No Warranty
HẾT BẢO HÀNH Out Of Warranty
| HỢP ĐỒNG BẢO TRÌ Service Contract |  |  | THUỘC ĐIỀU KHOẢN HỢP ĐỒNG No-charge
NGOÀI ĐIỀU KHOẢN HỢP ĐỒNG Charge
THUỘC MỘT PHẦN ĐIỀU KHOẢN HỢP ĐỒNG Particular Charge
ĐẠI DIỆN CÔNG TY OLYMPUS | [5 empty] | ĐẠI DIỆN KHÁCH HÀNG
Representative of Olympus Company | [5 empty] | Representative of Customer
NGƯỜI KIỂM TRA/ Inspected By
South Region Technical Service Engineer
NGƯỜI DUYỆT/ Approved By
South Region Technical Service Assistant Manager
OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD. (OVN)
Office in Hanoi: Unit 305- 307, 3rd Floor, V Tower, 649 Kim Ma Street, Ba Dinh District, HaNoi, Vietnam - Tel: (84-24) 3766 5831/32-Fax: (84-24) 3766 5833
Office in HoChiMinh City: 7th Floor, Viettel Building, 285 Cach Mang Thang Tam, Hoa Hung Ward, Ho Chi Minh City, Vietnam - Tel: (84-28) 3859 0028- Fax: (84-28) 3859 0029..
Warehouse and Warranty Center In HoChiMinh City: 2nd Floor, 86 Tan Đa, Cho Lon Ward, Ho Chi Minh City, Vietnam - Tel: (84-28) 3859 0028- Fax: (84-28) 3859 0029
Scope Structure
Endoscope Connector
Universal Cord
Control Section
Insertion Tube
Bending Section
Distal End

## GF-UCT180
BÁO CÁO GIÁM ĐỊNH KỸ THUẬT
TECHNICAL REPORT
SỐ/ No.: |  |  |  |  | NGÀY RA BÁO CÁO/ Issued Date:
KÍNH GỞI/ Report to: |  |  | Quý Khách Hàng/ Valued Customer
BỆNH VIỆN/ Name Of User:
KIỂU MÁY/ Name Of Model: |  |  |  |  | SỐ MÁY/ Serial No:
HƯ HỎNG TỪ KHÁCH HÀNG/ Customer Complaints:
1. 2.
ĐÁNH GIÁ TỪ OLYMPUS/ Olympus findings
Bộ Phận/ Section |  | Đánh Giá Hư Hỏng/ Confirmed Defects |  |  |  |  | Hình Ảnh/ Photos
Bộ phận uốn cong Bending Section |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Yêu cầu sửa chữa/ Repair required
Bộ phận điều khiển Control Unit |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Yêu cầu sửa chữa/ Repair required
Dây dẫn chung Universal Cord |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Khuyến nghị sửa chữa/ Repair recommended
GHI CHÚ/ Remark
NGUYÊN NHÂN HƯ HỎNG/ Reason Of Damages
- Các hư hỏng do hao mòn trong quá trình sử dụng/ The damages due to worn and tear during use.
GIẢI PHÁP/ Solutions
- Thay thế linh kiện như danh mục chuẩn đính kèm. Tuy nhiên nếu khách hàng không đồng ý sửa chữa lớn, chúng tôi chỉ thay thế linh kiện như danh mục khuyến nghị kèm theo. Chúng tôi chỉ bảo hành cho các linh kiện thay thế này.
Replace new spare parts as standard attachment list. However, if the customer does not agree to Major repair, we will replace the new spare parts as recommended attachment list. We only warrant replaced parts.
HÀNH ĐỘNG NGĂN NGỪA/ Recommended Actions To Prevent Sử dụng, bảo quản và vệ sinh tiệt trùng thiết bị theo đúng hướng dẫn của Olympus. Use, maintain and reprocess the equipment following Olympus’s instructions.
Sử dụng, bảo quản và vệ sinh tiệt trùng thiết bị theo đúng hướng dẫn của Olympus. Use, maintain and reprocess the equipment following Olympus’s instructions.
THÔNG TIN SỬA CHỮA/ Repair Information
| SỬA CHỮA THEO NHU CẦU On-demand Repair |  |  | TRONG BẢO HÀNH Under Warranty |  |  |  | BẢO HÀNH MUA MỚI Product Warranty
BẢO HÀNH SỬA CHỮA Repair Warranty
KHÔNG ĐƯỢC BẢO HÀNH No Warranty
HẾT BẢO HÀNH Out Of Warranty
| HỢP ĐỒNG BẢO TRÌ Service Contract |  |  | THUỘC ĐIỀU KHOẢN HỢP ĐỒNG No-charge
NGOÀI ĐIỀU KHOẢN HỢP ĐỒNG Charge
THUỘC MỘT PHẦN ĐIỀU KHOẢN HỢP ĐỒNG Particular Charge
ĐẠI DIỆN CÔNG TY OLYMPUS | [5 empty] | ĐẠI DIỆN KHÁCH HÀNG
Representative of Olympus Company | [5 empty] | Representative of Customer
NGƯỜI KIỂM TRA/ Inspected By
South Region Technical Service Engineer
NGƯỜI DUYỆT/ Approved By
South Region Technical Service Assistant Manager
OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD. (OVN)
Office in Hanoi: Unit 305- 307, 3rd Floor, V Tower, 649 Kim Ma Street, Ba Dinh District, HaNoi, Vietnam - Tel: (84-24) 3766 5831/32-Fax: (84-24) 3766 5833
Office in HoChiMinh City: 7th Floor, Viettel Building, 285 Cach Mang Thang Tam, Hoa Hung Ward, Ho Chi Minh City, Vietnam - Tel: (84-28) 3859 0028- Fax: (84-28) 3859 0029..
Warehouse and Warranty Center In HoChiMinh City: 2nd Floor, 86 Tan Đa, Cho Lon Ward, Ho Chi Minh City, Vietnam - Tel: (84-28) 3859 0028- Fax: (84-28) 3859 0029
Scope Structure
Endoscope Connector
Universal Cord
Control Section
Insertion Tube
Bending Section
Distal End

## CF-150, GIF-150,180
BÁO CÁO GIÁM ĐỊNH KỸ THUẬT
TECHNICAL REPORT
SỐ/ No.: |  |  |  |  | NGÀY RA BÁO CÁO/ Issued Date:
KÍNH GỞI/ Report to: |  |  | Quý Khách Hàng/ Valued Customer
BỆNH VIỆN/ Name Of User:
KIỂU MÁY/ Name Of Model: |  |  |  |  | SỐ MÁY/ Serial No:
HƯ HỎNG TỪ KHÁCH HÀNG/ Customer Complaints:
1. 2.
ĐÁNH GIÁ TỪ OLYMPUS/ Olympus findings
Bộ Phận/ Section |  | Đánh Giá Hư Hỏng/ Confirmed Defects |  |  |  |  | Hình Ảnh/ Photos
Bộ phận uốn cong Bending Section |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Yêu cầu sửa chữa/ Repair required
Bộ phận điều khiển Control Unit |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Yêu cầu sửa chữa/ Repair required
Dây dẫn chung Universal Cord |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Khuyến nghị sửa chữa/ Repair recommended
GHI CHÚ/ Remark
NGUYÊN NHÂN HƯ HỎNG/ Reason Of Damages
- Các hư hỏng do hao mòn trong quá trình sử dụng/ The damages due to worn and tear during use.
GIẢI PHÁP/ Solutions
- Thay thế linh kiện như danh mục chuẩn đính kèm. Tuy nhiên nếu khách hàng không đồng ý sửa chữa lớn, chúng tôi chỉ thay thế linh kiện như danh mục khuyến nghị kèm theo. Chúng tôi chỉ bảo hành cho các linh kiện thay thế này.
Replace new spare parts as standard attachment list. However, if the customer does not agree to Major repair, we will replace the new spare parts as recommended attachment list. We only warrant replaced parts.
HÀNH ĐỘNG NGĂN NGỪA/ Recommended Actions To Prevent Sử dụng, bảo quản và vệ sinh tiệt trùng thiết bị theo đúng hướng dẫn của Olympus. Use, maintain and reprocess the equipment following Olympus’s instructions.
Sử dụng, bảo quản và vệ sinh tiệt trùng thiết bị theo đúng hướng dẫn của Olympus. Use, maintain and reprocess the equipment following Olympus’s instructions.
THÔNG TIN SỬA CHỮA/ Repair Information
| SỬA CHỮA THEO NHU CẦU On-demand Repair |  |  | TRONG BẢO HÀNH Under Warranty |  |  |  | BẢO HÀNH MUA MỚI Product Warranty
BẢO HÀNH SỬA CHỮA Repair Warranty
KHÔNG ĐƯỢC BẢO HÀNH No Warranty
HẾT BẢO HÀNH Out Of Warranty
| HỢP ĐỒNG BẢO TRÌ Service Contract |  |  | THUỘC ĐIỀU KHOẢN HỢP ĐỒNG No-charge
NGOÀI ĐIỀU KHOẢN HỢP ĐỒNG Charge
THUỘC MỘT PHẦN ĐIỀU KHOẢN HỢP ĐỒNG Particular Charge
ĐẠI DIỆN CÔNG TY OLYMPUS | [5 empty] | ĐẠI DIỆN KHÁCH HÀNG
Representative of Olympus Company | [5 empty] | Representative of Customer
NGƯỜI KIỂM TRA/ Inspected By
South Region Technical Service Engineer
NGƯỜI DUYỆT/ Approved By
South Region Technical Service Assistant Manager
OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD. (OVN)
Office and Service Center in Hanoi: Unit 305- 307, 3rd Floor, V Tower, 649 Kim Ma Street, Giang Vo Ward, HaNoi City, Vietnam - Tel: (84-24) 3766 5831/32/34
Office In HoChiMinh City: 7th Floor, Viettel Building, 285 Cach Mang Thang Tam, Hoa Hung Ward, HoChiMinh City, Vietnam - Tel: (84-28) 7301 9919
Service Center In HoChiMinh City: 2nd Floor, 86 Tan Da, Cho Lon Ward, HoChiMinh City, Vietnam - Tel: (84-28) 7301 9919
Scope Structure
Endoscope Connector
Universal Cord
Control Section
Insertion Tube
Bending Section
Distal End

## TJF-190
BÁO CÁO GIÁM ĐỊNH KỸ THUẬT
TECHNICAL REPORT
SỐ/ No.: |  |  |  |  | NGÀY RA BÁO CÁO/ Issued Date:
KÍNH GỞI/ Report to: |  |  | Quý Khách Hàng/ Valued Customer
BỆNH VIỆN/ Name Of User:
KIỂU MÁY/ Name Of Model: |  |  |  |  | SỐ MÁY/ Serial No:
HƯ HỎNG TỪ KHÁCH HÀNG/ Customer Complaints:
1. 2.
ĐÁNH GIÁ TỪ OLYMPUS/ Olympus findings
Bộ Phận/ Section |  | Đánh Giá Hư Hỏng/ Confirmed Defects |  |  |  |  | Hình Ảnh/ Photos
Bộ phận uốn cong Bending Section |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Yêu cầu sửa chữa/ Repair required
Bộ phận điều khiển Control Unit |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Yêu cầu sửa chữa/ Repair required
Dây dẫn chung Universal Cord |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Khuyến nghị sửa chữa/ Repair recommended
GHI CHÚ/ Remark
NGUYÊN NHÂN HƯ HỎNG/ Reason Of Damages
- Các hư hỏng do hao mòn trong quá trình sử dụng/ The damages due to worn and tear during use.
GIẢI PHÁP/ Solutions
- Thay thế linh kiện như danh mục chuẩn đính kèm. Tuy nhiên nếu khách hàng không đồng ý sửa chữa lớn, chúng tôi chỉ thay thế linh kiện như danh mục khuyến nghị kèm theo. Chúng tôi chỉ bảo hành cho các linh kiện thay thế này.
Replace new spare parts as standard attachment list. However, if the customer does not agree to Major repair, we will replace the new spare parts as recommended attachment list. We only warrant replaced parts.
HÀNH ĐỘNG NGĂN NGỪA/ Recommended Actions To Prevent Sử dụng, bảo quản và vệ sinh tiệt trùng thiết bị theo đúng hướng dẫn của Olympus. Use, maintain and reprocess the equipment following Olympus’s instructions.
Sử dụng, bảo quản và vệ sinh tiệt trùng thiết bị theo đúng hướng dẫn của Olympus. Use, maintain and reprocess the equipment following Olympus’s instructions.
THÔNG TIN SỬA CHỮA/ Repair Information
| SỬA CHỮA THEO NHU CẦU On-demand Repair |  |  | TRONG BẢO HÀNH Under Warranty |  |  |  | BẢO HÀNH MUA MỚI Product Warranty
BẢO HÀNH SỬA CHỮA Repair Warranty
KHÔNG ĐƯỢC BẢO HÀNH No Warranty
HẾT BẢO HÀNH Out Of Warranty
| HỢP ĐỒNG BẢO TRÌ Service Contract |  |  | THUỘC ĐIỀU KHOẢN HỢP ĐỒNG No-charge
NGOÀI ĐIỀU KHOẢN HỢP ĐỒNG Charge
THUỘC MỘT PHẦN ĐIỀU KHOẢN HỢP ĐỒNG Particular Charge
ĐẠI DIỆN CÔNG TY OLYMPUS | [5 empty] | ĐẠI DIỆN KHÁCH HÀNG
Representative of Olympus Company | [5 empty] | Representative of Customer
NGƯỜI KIỂM TRA/ Inspected By
South Region Technical Service Engineer
NGƯỜI DUYỆT/ Approved By
South Region Technical Service Assistant Manager
OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD. (OVN)
Office in Hanoi: Unit 305- 307, 3rd Floor, V Tower, 649 Kim Ma Street, Ba Dinh District, HaNoi, Vietnam - Tel: (84-24) 3766 5831/32-Fax: (84-24) 3766 5833
Office in HoChiMinh City: 7th Floor, Viettel Building, 285 Cach Mang Thang Tam, Hoa Hung Ward, Ho Chi Minh City, Vietnam - Tel: (84-28) 3859 0028- Fax: (84-28) 3859 0029..
Warehouse and Warranty Center In HoChiMinh City: 2nd Floor, 86 Tan Đa, Cho Lon Ward, Ho Chi Minh City, Vietnam - Tel: (84-28) 3859 0028- Fax: (84-28) 3859 0029
Scope Structure
Endoscope Connector
Universal Cord
Control Section
Insertion Tube
Bending Section
Distal End

## TJF-180
BÁO CÁO GIÁM ĐỊNH KỸ THUẬT
TECHNICAL REPORT
SỐ/ No.: |  |  |  |  | NGÀY RA BÁO CÁO/ Issued Date:
KÍNH GỞI/ Report to: |  |  | Quý Khách Hàng/ Valued Customer
BỆNH VIỆN/ Name Of User:
KIỂU MÁY/ Name Of Model: |  |  |  |  | SỐ MÁY/ Serial No:
HƯ HỎNG TỪ KHÁCH HÀNG/ Customer Complaints:
1. 2.
ĐÁNH GIÁ TỪ OLYMPUS/ Olympus findings
Bộ Phận/ Section |  | Đánh Giá Hư Hỏng/ Confirmed Defects |  |  |  |  | Hình Ảnh/ Photos
Bộ phận uốn cong Bending Section |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Yêu cầu sửa chữa/ Repair required
Bộ phận điều khiển Control Unit |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Yêu cầu sửa chữa/ Repair required
Dây dẫn chung Universal Cord |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Khuyến nghị sửa chữa/ Repair recommended
GHI CHÚ/ Remark
NGUYÊN NHÂN HƯ HỎNG/ Reason Of Damages
- Các hư hỏng do hao mòn trong quá trình sử dụng/ The damages due to worn and tear during use.
GIẢI PHÁP/ Solutions
- Thay thế linh kiện như danh mục chuẩn đính kèm. Tuy nhiên nếu khách hàng không đồng ý sửa chữa lớn, chúng tôi chỉ thay thế linh kiện như danh mục khuyến nghị kèm theo. Chúng tôi chỉ bảo hành cho các linh kiện thay thế này.
Replace new spare parts as standard attachment list. However, if the customer does not agree to Major repair, we will replace the new spare parts as recommended attachment list. We only warrant replaced parts.
HÀNH ĐỘNG NGĂN NGỪA/ Recommended Actions To Prevent Sử dụng, bảo quản và vệ sinh tiệt trùng thiết bị theo đúng hướng dẫn của Olympus. Use, maintain and reprocess the equipment following Olympus’s instructions.
Sử dụng, bảo quản và vệ sinh tiệt trùng thiết bị theo đúng hướng dẫn của Olympus. Use, maintain and reprocess the equipment following Olympus’s instructions.
THÔNG TIN SỬA CHỮA/ Repair Information
| SỬA CHỮA THEO NHU CẦU On-demand Repair |  |  | TRONG BẢO HÀNH Under Warranty |  |  |  | BẢO HÀNH MUA MỚI Product Warranty
BẢO HÀNH SỬA CHỮA Repair Warranty
KHÔNG ĐƯỢC BẢO HÀNH No Warranty
HẾT BẢO HÀNH Out Of Warranty
| HỢP ĐỒNG BẢO TRÌ Service Contract |  |  | THUỘC ĐIỀU KHOẢN HỢP ĐỒNG No-charge
NGOÀI ĐIỀU KHOẢN HỢP ĐỒNG Charge
THUỘC MỘT PHẦN ĐIỀU KHOẢN HỢP ĐỒNG Particular Charge
ĐẠI DIỆN CÔNG TY OLYMPUS | [5 empty] | ĐẠI DIỆN KHÁCH HÀNG
Representative of Olympus Company | [5 empty] | Representative of Customer
NGƯỜI KIỂM TRA/ Inspected By
South Region Technical Service Engineer
NGƯỜI DUYỆT/ Approved By
South Region Technical Service Assistant Manager
OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD. (OVN)
Office in Hanoi: Unit 305- 307, 3rd Floor, V Tower, 649 Kim Ma Street, Ba Dinh District, HaNoi, Vietnam - Tel: (84-24) 3766 5831/32-Fax: (84-24) 3766 5833
Office in HoChiMinh City: 7th Floor, Viettel Building, 285 Cach Mang Thang Tam, Hoa Hung Ward, Ho Chi Minh City, Vietnam - Tel: (84-28) 3859 0028- Fax: (84-28) 3859 0029..
Warehouse and Warranty Center In HoChiMinh City: 2nd Floor, 86 Tan Đa, Cho Lon Ward, Ho Chi Minh City, Vietnam - Tel: (84-28) 3859 0028- Fax: (84-28) 3859 0029
Scope Structure
Endoscope Connector
Universal Cord
Control Section
Insertion Tube
Bending Section
Distal End

## TJF-Q170V
BÁO CÁO GIÁM ĐỊNH KỸ THUẬT
TECHNICAL REPORT
SỐ/ No.: |  |  |  |  | NGÀY RA BÁO CÁO/ Issued Date:
KÍNH GỞI/ Report to: |  |  | Quý Khách Hàng/ Valued Customer
BỆNH VIỆN/ Name Of User:
KIỂU MÁY/ Name Of Model: |  |  |  |  | SỐ MÁY/ Serial No:
HƯ HỎNG TỪ KHÁCH HÀNG/ Customer Complaints:
1. 2.
ĐÁNH GIÁ TỪ OLYMPUS/ Olympus findings
Bộ Phận/ Section |  | Đánh Giá Hư Hỏng/ Confirmed Defects |  |  |  |  | Hình Ảnh/ Photos
Bộ phận uốn cong Bending Section |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Yêu cầu sửa chữa/ Repair required
Bộ phận điều khiển Control Unit |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Yêu cầu sửa chữa/ Repair required
Dây dẫn chung Universal Cord |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Khuyến nghị sửa chữa/ Repair recommended
GHI CHÚ/ Remark
NGUYÊN NHÂN HƯ HỎNG/ Reason Of Damages
- Các hư hỏng do hao mòn trong quá trình sử dụng/ The damages due to worn and tear during use.
GIẢI PHÁP/ Solutions
- Thay thế linh kiện như danh mục chuẩn đính kèm. Tuy nhiên nếu khách hàng không đồng ý sửa chữa lớn, chúng tôi chỉ thay thế linh kiện như danh mục khuyến nghị kèm theo. Chúng tôi chỉ bảo hành cho các linh kiện thay thế này.
Replace new spare parts as standard attachment list. However, if the customer does not agree to Major repair, we will replace the new spare parts as recommended attachment list. We only warrant replaced parts.
HÀNH ĐỘNG NGĂN NGỪA/ Recommended Actions To Prevent Sử dụng, bảo quản và vệ sinh tiệt trùng thiết bị theo đúng hướng dẫn của Olympus. Use, maintain and reprocess the equipment following Olympus’s instructions.
Sử dụng, bảo quản và vệ sinh tiệt trùng thiết bị theo đúng hướng dẫn của Olympus. Use, maintain and reprocess the equipment following Olympus’s instructions.
THÔNG TIN SỬA CHỮA/ Repair Information
| SỬA CHỮA THEO NHU CẦU On-demand Repair |  |  | TRONG BẢO HÀNH Under Warranty |  |  |  | BẢO HÀNH MUA MỚI Product Warranty
BẢO HÀNH SỬA CHỮA Repair Warranty
KHÔNG ĐƯỢC BẢO HÀNH No Warranty
HẾT BẢO HÀNH Out Of Warranty
| HỢP ĐỒNG BẢO TRÌ Service Contract |  |  | THUỘC ĐIỀU KHOẢN HỢP ĐỒNG No-charge
NGOÀI ĐIỀU KHOẢN HỢP ĐỒNG Charge
THUỘC MỘT PHẦN ĐIỀU KHOẢN HỢP ĐỒNG Particular Charge
ĐẠI DIỆN CÔNG TY OLYMPUS | [5 empty] | ĐẠI DIỆN KHÁCH HÀNG
Representative of Olympus Company | [5 empty] | Representative of Customer
NGƯỜI KIỂM TRA/ Inspected By
South Region Technical Service Engineer
NGƯỜI DUYỆT/ Approved By
South Region Technical Service Assistant Manager
OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD. (OVN)
Office in Hanoi: Unit 305- 307, 3rd Floor, V Tower, 649 Kim Ma Street, Ba Dinh District, HaNoi, Vietnam - Tel: (84-24) 3766 5831/32-Fax: (84-24) 3766 5833
Office in HoChiMinh City: 7th Floor, Viettel Building, 285 Cach Mang Thang Tam, Hoa Hung Ward, Ho Chi Minh City, Vietnam - Tel: (84-28) 3859 0028- Fax: (84-28) 3859 0029..
Warehouse and Warranty Center In HoChiMinh City: 2nd Floor, 86 Tan Đa, Cho Lon Ward, Ho Chi Minh City, Vietnam - Tel: (84-28) 3859 0028- Fax: (84-28) 3859 0029
Scope Structure
Universal Cord
Control Section
Insertion Tube
Bending Section
Distal End
LG Connector
Video Connector

## TJF-160
BÁO CÁO GIÁM ĐỊNH KỸ THUẬT
TECHNICAL REPORT
SỐ/ No.: |  |  |  |  | NGÀY RA BÁO CÁO/ Issued Date:
KÍNH GỞI/ Report to: |  |  | Quý Khách Hàng/ Valued Customer
BỆNH VIỆN/ Name Of User:
KIỂU MÁY/ Name Of Model: |  |  |  |  | SỐ MÁY/ Serial No:
HƯ HỎNG TỪ KHÁCH HÀNG/ Customer Complaints:
1. 2.
ĐÁNH GIÁ TỪ OLYMPUS/ Olympus findings
Bộ Phận/ Section |  | Đánh Giá Hư Hỏng/ Confirmed Defects |  |  |  |  | Hình Ảnh/ Photos
Bộ phận uốn cong Bending Section |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Yêu cầu sửa chữa/ Repair required
Bộ phận điều khiển Control Unit |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Yêu cầu sửa chữa/ Repair required
Dây dẫn chung Universal Cord |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Khuyến nghị sửa chữa/ Repair recommended
GHI CHÚ/ Remark
NGUYÊN NHÂN HƯ HỎNG/ Reason Of Damages
- Các hư hỏng do hao mòn trong quá trình sử dụng/ The damages due to worn and tear during use.
GIẢI PHÁP/ Solutions
- Thay thế linh kiện như danh mục chuẩn đính kèm. Tuy nhiên nếu khách hàng không đồng ý sửa chữa lớn, chúng tôi chỉ thay thế linh kiện như danh mục khuyến nghị kèm theo. Chúng tôi chỉ bảo hành cho các linh kiện thay thế này.
Replace new spare parts as standard attachment list. However, if the customer does not agree to Major repair, we will replace the new spare parts as recommended attachment list. We only warrant replaced parts.
HÀNH ĐỘNG NGĂN NGỪA/ Recommended Actions To Prevent Sử dụng, bảo quản và vệ sinh tiệt trùng thiết bị theo đúng hướng dẫn của Olympus. Use, maintain and reprocess the equipment following Olympus’s instructions.
Sử dụng, bảo quản và vệ sinh tiệt trùng thiết bị theo đúng hướng dẫn của Olympus. Use, maintain and reprocess the equipment following Olympus’s instructions.
THÔNG TIN SỬA CHỮA/ Repair Information
| SỬA CHỮA THEO NHU CẦU On-demand Repair |  |  | TRONG BẢO HÀNH Under Warranty |  |  |  | BẢO HÀNH MUA MỚI Product Warranty
BẢO HÀNH SỬA CHỮA Repair Warranty
KHÔNG ĐƯỢC BẢO HÀNH No Warranty
HẾT BẢO HÀNH Out Of Warranty
| HỢP ĐỒNG BẢO TRÌ Service Contract |  |  | THUỘC ĐIỀU KHOẢN HỢP ĐỒNG No-charge
NGOÀI ĐIỀU KHOẢN HỢP ĐỒNG Charge
THUỘC MỘT PHẦN ĐIỀU KHOẢN HỢP ĐỒNG Particular Charge
ĐẠI DIỆN CÔNG TY OLYMPUS | [5 empty] | ĐẠI DIỆN KHÁCH HÀNG
Representative of Olympus Company | [5 empty] | Representative of Customer
NGƯỜI KIỂM TRA/ Inspected By
South Region Technical Service Engineer
NGƯỜI DUYỆT/ Approved By
South Region Technical Service Assistant Manager
OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD. (OVN)
Office in Hanoi: Unit 305- 307, 3rd Floor, V Tower, 649 Kim Ma Street, Ba Dinh District, HaNoi, Vietnam - Tel: (84-24) 3766 5831/32-Fax: (84-24) 3766 5833
Office in HoChiMinh City: 7th Floor, Viettel Building, 285 Cach Mang Thang Tam, Hoa Hung Ward, Ho Chi Minh City, Vietnam - Tel: (84-28) 3859 0028- Fax: (84-28) 3859 0029..
Warehouse and Warranty Center In HoChiMinh City: 2nd Floor, 86 Tan Đa, Cho Lon Ward, Ho Chi Minh City, Vietnam - Tel: (84-28) 3859 0028- Fax: (84-28) 3859 0029
Scope Structure
Endoscope Connector
Universal Cord
Control Section
Insertion Tube
Bending Section
Distal End

## CF, GIF-LV1
BÁO CÁO GIÁM ĐỊNH KỸ THUẬT
TECHNICAL REPORT
SỐ/ No.: |  |  |  |  | NGÀY RA BÁO CÁO/ Issued Date:
KÍNH GỞI/ Report to: |  |  | Quý Khách Hàng/ Valued Customer
BỆNH VIỆN/ Name Of User:
KIỂU MÁY/ Name Of Model: |  |  |  |  | SỐ MÁY/ Serial No:
HƯ HỎNG TỪ KHÁCH HÀNG/ Customer Complaints:
1. 2.
ĐÁNH GIÁ TỪ OLYMPUS/ Olympus findings
Bộ Phận/ Section |  | Đánh Giá Hư Hỏng/ Confirmed Defects |  |  |  |  | Hình Ảnh/ Photos
Bộ phận uốn cong Bending Section |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Yêu cầu sửa chữa/ Repair required
Bộ phận điều khiển Control Unit |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Yêu cầu sửa chữa/ Repair required
Dây dẫn chung Universal Cord |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Khuyến nghị sửa chữa/ Repair recommended
GHI CHÚ/ Remark
NGUYÊN NHÂN HƯ HỎNG/ Reason Of Damages
- Các hư hỏng do hao mòn trong quá trình sử dụng/ The damages due to worn and tear during use.
GIẢI PHÁP/ Solutions
- Thay thế linh kiện như danh mục chuẩn đính kèm. Tuy nhiên nếu khách hàng không đồng ý sửa chữa lớn, chúng tôi chỉ thay thế linh kiện như danh mục khuyến nghị kèm theo. Chúng tôi chỉ bảo hành cho các linh kiện thay thế này.
Replace new spare parts as standard attachment list. However, if the customer does not agree to Major repair, we will replace the new spare parts as recommended attachment list. We only warrant replaced parts.
HÀNH ĐỘNG NGĂN NGỪA/ Recommended Actions To Prevent Sử dụng, bảo quản và vệ sinh tiệt trùng thiết bị theo đúng hướng dẫn của Olympus. Use, maintain and reprocess the equipment following Olympus’s instructions.
Sử dụng, bảo quản và vệ sinh tiệt trùng thiết bị theo đúng hướng dẫn của Olympus. Use, maintain and reprocess the equipment following Olympus’s instructions.
THÔNG TIN SỬA CHỮA/ Repair Information
| SỬA CHỮA THEO NHU CẦU On-demand Repair |  |  | TRONG BẢO HÀNH Under Warranty |  |  |  | BẢO HÀNH MUA MỚI Product Warranty
BẢO HÀNH SỬA CHỮA Repair Warranty
KHÔNG ĐƯỢC BẢO HÀNH No Warranty
HẾT BẢO HÀNH Out Of Warranty
| HỢP ĐỒNG BẢO TRÌ Service Contract |  |  | THUỘC ĐIỀU KHOẢN HỢP ĐỒNG No-charge
NGOÀI ĐIỀU KHOẢN HỢP ĐỒNG Charge
THUỘC MỘT PHẦN ĐIỀU KHOẢN HỢP ĐỒNG Particular Charge
ĐẠI DIỆN CÔNG TY OLYMPUS | [5 empty] | ĐẠI DIỆN KHÁCH HÀNG
Representative of Olympus Company | [5 empty] | Representative of Customer
NGƯỜI KIỂM TRA/ Inspected By
South Region Technical Service Engineer
NGƯỜI DUYỆT/ Approved By
South Region Technical Service Assistant Manager
OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD. (OVN)
Office in Hanoi: Unit 305- 307, 3rd Floor, V Tower, 649 Kim Ma Street, Ba Dinh District, HaNoi, Vietnam - Tel: (84-24) 3766 5831/32-Fax: (84-24) 3766 5833
Office in HoChiMinh City: 7th Floor, Viettel Building, 285 Cach Mang Thang Tam, Hoa Hung Ward, Ho Chi Minh City, Vietnam - Tel: (84-28) 3859 0028- Fax: (84-28) 3859 0029..
Warehouse and Warranty Center In HoChiMinh City: 2nd Floor, 86 Tan Đa, Cho Lon Ward, Ho Chi Minh City, Vietnam - Tel: (84-28) 3859 0028- Fax: (84-28) 3859 0029
Scope Structure
Endoscope Connector
Universal Cord
Control Section
Insertion Tube
Bending Section
Distal End

## GIF-V70
BÁO CÁO GIÁM ĐỊNH KỸ THUẬT
TECHNICAL REPORT
SỐ/ No.: |  |  |  |  | NGÀY RA BÁO CÁO/ Issued Date:
KÍNH GỞI/ Report to: |  |  | Quý Khách Hàng/ Valued Customer
BỆNH VIỆN/ Name Of User:
KIỂU MÁY/ Name Of Model: |  |  |  |  | SỐ MÁY/ Serial No:
HƯ HỎNG TỪ KHÁCH HÀNG/ Customer Complaints:
1. 2.
ĐÁNH GIÁ TỪ OLYMPUS/ Olympus findings
Bộ Phận/ Section |  | Đánh Giá Hư Hỏng/ Confirmed Defects |  |  |  |  | Hình Ảnh/ Photos
Bộ phận uốn cong Bending Section |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Yêu cầu sửa chữa/ Repair required
Bộ phận điều khiển Control Unit |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Yêu cầu sửa chữa/ Repair required
Dây dẫn chung Universal Cord |  | Hiện Trạng/ Phenomenon
Đánh giá/ Judgement
Khuyến nghị sửa chữa/ Repair recommended
GHI CHÚ/ Remark
NGUYÊN NHÂN HƯ HỎNG/ Reason Of Damages
- Các hư hỏng do hao mòn trong quá trình sử dụng/ The damages due to worn and tear during use.
GIẢI PHÁP/ Solutions
- Thay thế linh kiện như danh mục chuẩn đính kèm. Tuy nhiên nếu khách hàng không đồng ý sửa chữa lớn, chúng tôi chỉ thay thế linh kiện như danh mục khuyến nghị kèm theo. Chúng tôi chỉ bảo hành cho các linh kiện thay thế này.
Replace new spare parts as standard attachment list. However, if the customer does not agree to Major repair, we will replace the new spare parts as recommended attachment list. We only warrant replaced parts.
HÀNH ĐỘNG NGĂN NGỪA/ Recommended Actions To Prevent Sử dụng, bảo quản và vệ sinh tiệt trùng thiết bị theo đúng hướng dẫn của Olympus. Use, maintain and reprocess the equipment following Olympus’s instructions.
Sử dụng, bảo quản và vệ sinh tiệt trùng thiết bị theo đúng hướng dẫn của Olympus. Use, maintain and reprocess the equipment following Olympus’s instructions.
THÔNG TIN SỬA CHỮA/ Repair Information
| SỬA CHỮA THEO NHU CẦU On-demand Repair |  |  | TRONG BẢO HÀNH Under Warranty |  |  |  | BẢO HÀNH MUA MỚI Product Warranty
BẢO HÀNH SỬA CHỮA Repair Warranty
KHÔNG ĐƯỢC BẢO HÀNH No Warranty
HẾT BẢO HÀNH Out Of Warranty
| HỢP ĐỒNG BẢO TRÌ Service Contract |  |  | THUỘC ĐIỀU KHOẢN HỢP ĐỒNG No-charge
NGOÀI ĐIỀU KHOẢN HỢP ĐỒNG Charge
THUỘC MỘT PHẦN ĐIỀU KHOẢN HỢP ĐỒNG Particular Charge
ĐẠI DIỆN CÔNG TY OLYMPUS | [5 empty] | ĐẠI DIỆN KHÁCH HÀNG
Representative of Olympus Company | [5 empty] | Representative of Customer
NGƯỜI KIỂM TRA/ Inspected By
South Region Technical Service Engineer
NGƯỜI DUYỆT/ Approved By
South Region Technical Service Assistant Manager
OLYMPUS MEDICAL SYSTEMS VIETNAM CO.,LTD. (OVN)
Office in Hanoi: Unit 305- 307, 3rd Floor, V Tower, 649 Kim Ma Street, Ba Dinh District, HaNoi, Vietnam - Tel: (84-24) 3766 5831/32-Fax: (84-24) 3766 5833.
Office in HoChiMinh City: 7th Floor, Viettel Building, 285 Cach Mang Thang Tam, Hoa Hung Ward, Ho Chi Minh City, Vietnam - Tel: (84-28) 3859 0028- Fax: (84-28) 3859 0029. - Tel: (84-28) 3859 0028- Fax: (84-28) 3859 0029.
Warehouse and Warranty Center In HoChiMinh City: 2nd Floor, 86 Tan Đa, Cho Lon Ward, Ho Chi Minh City, Vietnam - Tel: (84-28) 3859 0028- Fax: (84-28) 3859 0029.
Scope Structure
Endoscope Connector
Universal Cord
Control Section
Insertion Tube
Bending Section
Distal End